## Temperature
MOLBOX RESISTANCE/TEMPERATURE CALIBRATION
BOX SN:
COMBIND TEMPERATURE TOLERANCE: |  |  | 0.05
MOLBLOC SIM UNCERTANITY @ K=1 |  |  | 7.5e-05
TEMPERATURE (RESISTANCE) MEASUREMENTS | [6 empty] | TEMPERATURE (RESISTANCE) MEASUREMENTS
| Reference | Measurement | Measurement | Specification (+) | Specification (-) | Pass/Fail |  | Reference | Measurement | Measurement | Specification (+) | Specification (-) | Pass/Fail
| Value | Result | Error |  |  |  |  | Value | Result | Error
AS FOUND CHANNEL A | [6 empty] | AS FOUND CHANNEL B
°C T1 | 0 | -0.05 | -0.045 | 0.05 | -0.05 | Pass | °C T1 | 0 | -0.04 | -0.04 | 0.05 | -0.05 | Pass
T2 | 0 | -0.04 |  |  |  |  | T2 | 0 | -0.04
°C T1 | 25.67 | 25.65 | -0.025 | 0.05 | -0.05 | Pass | °C T1 | 25.67 | 25.64 | -0.035 | 0.05 | -0.05 | Pass
T2 | 25.67 | 25.64 |  |  |  |  | T2 | 25.67 | 25.63
°C T1 | 38.5 | 38.45 | -0.05 | 0.05 | -0.05 | Pass | °C T1 | 38.5 | 38.44 | -0.065 | 0.05 | -0.05 | Fail
T2 | 38.5 | 38.45 |  |  |  |  | T2 | 38.5 | 38.43
Temp Adder | 0 | [5 empty] | Temp Adder | 0
Temp Multiplier | 1 | [5 empty] | Temp Multiplier | 1
| Reference | Measurement | Measurement | Specification | Specification (-) | Pass/Fail |  | Reference | Measurement | Measurement | Specification | Specification (-) | Pass/Fail
| Value | Result | Error |  |  |  |  | Value | Result | Error
AS FOUND CHANNEL A WITHOUT AS FOUND TA AND TM | [6 empty] | AS FOUND CHANNEL B WITHOUT AS FOUND TA AND TM
°C T1 | 0 | -0.05 | -0.045 | 0.05 | -0.05 | Pass | °C T1 | 0 | -0.04 | -0.04 | 0.05 | -0.05 | Pass
T2 | 0 | -0.04 |  |  |  |  | T2 | 0 | -0.04
°C T1 | 25.67 | 25.65 | -0.025 | 0.05 | -0.05 | Pass | °C T1 | 25.67 | 25.64 | -0.035 | 0.05 | -0.05 | Pass
T2 | 25.67 | 25.64 |  |  |  |  | T2 | 25.67 | 25.63
°C T1 | 38.5 | 38.45 | -0.05 | 0.05 | -0.05 | Pass | °C T1 | 38.5 | 38.44 | -0.065 | 0.05 | -0.05 | Fail
T2 | 38.5 | 38.45 |  |  |  |  | T2 | 38.5 | 38.43
New Temp Adder | 0.040012 | [5 empty] | New Temp Adder | 0.035387
New Temp Multiplier | 1 | [5 empty] | New Temp Multiplier | 1.0005
| Reference | Measurement | Measurement | Specification | Specification (-) | Pass/Fail |  | Reference | Measurement | Measurement | Specification | Specification (-) | Pass/Fail
| Value | Result | Error |  |  |  |  | Value | Result | Error
AS LEFT CHANNEL A (PREDICTED) WITH AS LEFT TA AND TM | [6 empty] | AS LEFT CHANNEL B (PREDICTED) WITH AS LEFT TA AND TM
°C T1 | 0 | -0.0099875 | -0.0049875 | 0.05 | -0.05 | Pass | °C T1 | 0 | 1.2523e-05 | 1.2523e-05 | 0.05 | -0.05 | Pass
T2 | 0 | 1.2523e-05 |  |  |  |  | T2 | 0 | 1.2523e-05
°C T1 | 25.67 | 25.69 | 0.014997 | 0.05 | -0.05 | Pass | °C T1 | 25.67 | 25.68 | 0.0049975 | 0.05 | -0.05 | Pass
T2 | 25.67 | 25.68 |  |  |  |  | T2 | 25.67 | 25.67
°C T1 | 38.5 | 38.49 | -0.01001 | 0.05 | -0.05 | Pass | °C T1 | 38.5 | 38.48 | -0.02501 | 0.05 | -0.05 | Pass
T2 | 38.5 | 38.49 |  |  |  |  | T2 | 38.5 | 38.47
| Reference | Measurement | Measurement | Specification (+) | Specification (-) | Pass/Fail |  | Reference | Measurement | Measurement | Specification (+) | Specification (-) | Pass/Fail
| Value | Result | Error |  |  |  |  | Value | Result | Error
AS LEFT CHANNEL A | [6 empty] | AS LEFT CHANNEL B
°C T1 | 0 | 0.028 | 0.0295 | 0.05 | -0.05 | Pass | °C T1 | 0 | 0.037 | 0.031 | 0.05 | -0.05 | Pass
T2 | 0 | 0.031 |  |  |  |  | T2 | 0 | 0.025
°C T1 | 25.67 | 25.71 | 0.05 | 0.05 | -0.05 | Pass | °C T1 | 25.67 | 25.71 | 0.035 | 0.05 | -0.05 | Pass
T2 | 25.67 | 25.73 |  |  |  |  | T2 | 25.67 | 25.7
°C T1 | 38.5 | 38.5 | 0.015 | 0.05 | -0.05 | Pass | °C T1 | 38.5 | 38.5 | 0 | 0.05 | -0.05 | Pass
T2 | 38.5 | 38.53 |  |  |  |  | T2 | 38.5 | 38.5
Temp Adder | 0.040012 | [5 empty] | Temp Adder | 0.035387
Temp Multiplier | 1 | [5 empty] | Temp Multiplier | 1.0005

## Pressure
molbox1+ Pressure Calibration
model: | molbox1+ A700K
SN: | 1234
Specifications/Tolerances
As Found | %Rdg | Threshold %FS | 1-Year Drift | FS |  | Change 1-Year Drift to 0% if you regularly AutoZero the transducers
| 0.01% | 0.003% | 0.005% | 700 | kPa
As Left | %Rdg | Threshold %FS
| 0.008% | 0.0024%
AS FOUND DATA FOR UPSTREAM (HIGH) TRANSDUCER | [5 empty] | AS FOUND DATA FOR DOWNSTREAM (LOW) TRANSDUCER
Reference | Test | Error | As Found Spec(+) | As Found Spec(-) | Pass/Fail | Reference | Test | Error | As Found Spec(+) | As Found Spec(-) | Pass/Fail
kPa | kPa | kPa | kPa | kPa |  | kPa | kPa | kPa | kPa | kPa
19.8531 | 19.877 | 0.0239 | 0.056 | -0.056 | Pass | 19.8531 | 19.849 | -0.0041 | 0.056 | -0.056 | Pass
99.9873 | 100.007 | 0.0197 | 0.056 | -0.056 | Pass | 99.9873 | 99.982 | -0.0053 | 0.056 | -0.056 | Pass
200.7614 | 200.78 | 0.0186 | 0.056 | -0.056 | Pass | 200.7614 | 200.755 | -0.0064 | 0.056 | -0.056 | Pass
299.1284 | 299.15 | 0.0216 | 0.064913 | -0.064913 | Pass | 299.1284 | 299.121 | -0.0074 | 0.064913 | -0.064913 | Pass
449.8313 | 449.856 | 0.0247 | 0.079983 | -0.079983 | Pass | 449.8313 | 449.824 | -0.0073 | 0.079983 | -0.079983 | Pass
600.266 | 600.296 | 0.03 | 0.095027 | -0.095027 | Pass | 600.266 | 600.261 | -0.005 | 0.095027 | -0.095027 | Pass
449.8311 | 449.858 | 0.0269 | 0.079983 | -0.079983 | Pass | 449.8311 | 449.826 | -0.0051 | 0.079983 | -0.079983 | Pass
299.1281 | 299.153 | 0.0249 | 0.064913 | -0.064913 | Pass | 299.1281 | 299.123 | -0.0051 | 0.064913 | -0.064913 | Pass
200.761 | 200.784 | 0.023 | 0.056 | -0.056 | Pass | 200.761 | 200.757 | -0.004 | 0.056 | -0.056 | Pass
99.987 | 100.011 | 0.024 | 0.056 | -0.056 | Pass | 99.987 | 99.983 | -0.004 | 0.056 | -0.056 | Pass
19.8529 | 19.878 | 0.0251 | 0.056 | -0.056 | Pass | 19.8529 | 19.85 | -0.0029 | 0.056 | -0.056 | Pass
As Found Specification: | ±(0.010% of reading or 0.0030% of FS, whichever is greater), plus 0.005% of FS for one year drift without AutoZero |  |  |  |  | As Found Specification: | ±(0.010% of reading or 0.0030% of FS, whichever is greater), plus 0.005% of FS for one year drift without AutoZero
Pressure Adder (PA): | 14.9 | Pa |  |  |  | Pressure Adder (PA): | 133.4 | Pa
Pressure Multiplier (PM): | 1 |  |  |  |  | Pressure Multiplier (PM): | 1.0001
AutoZ (zOffset): | 0 | Pa |  |  |  | AutoZ (zOffset): | 0 | Pa
AS FOUND DATA FOR UPSTREAM (HIGH) TRANSDUCER WITHOUT AS FOUND PA AND PM | [5 empty] | AS FOUND DATA FOR DOWNSTREAM (LOW) TRANSDUCER WITHOUT AS FOUND PA AND PM
Reference | Test | Error |  |  |  | Reference | Test | Error
kPa | kPa | kPa |  |  |  | kPa | kPa | kPa
19.8531 | 19.8619 | 0.0088212 |  |  |  | 19.8531 | 19.7139 | -0.13916
99.9873 | 99.9912 | 0.0039001 |  |  |  | 99.9873 | 99.8402 | -0.14709
200.7614 | 200.7633 | 0.0018931 |  |  |  | 200.7614 | 200.6047 | -0.15665
299.1284 | 299.1324 | 0.0040078 |  |  |  | 299.1284 | 298.9625 | -0.16591
449.8313 | 449.8371 | 0.0057515 |  |  |  | 449.8313 | 449.6528 | -0.17847
600.266 | 600.2757 | 0.0096975 |  |  |  | 600.266 | 600.0772 | -0.18881
449.8311 | 449.8391 | 0.0079514 |  |  |  | 449.8311 | 449.6548 | -0.17627
299.1281 | 299.1354 | 0.0073078 |  |  |  | 299.1281 | 298.9645 | -0.16361
200.761 | 200.7673 | 0.0062931 |  |  |  | 200.761 | 200.6067 | -0.15425
99.987 | 99.9952 | 0.0082 |  |  |  | 99.987 | 99.8412 | -0.14579
19.8529 | 19.8629 | 0.010021 |  |  |  | 19.8529 | 19.7149 | -0.13796
New Pressure Adder (PA): | -6.5709 | Pa |  |  |  | New Pressure Adder (PA): | 137.5999 | Pa
New Pressure Multiplier (PM): | 1 |  |  |  |  | New Pressure Multiplier (PM): | 1.0001
AutoZ (zOffset): | 0 | Pa |  |  |  | AutoZ (zOffset): | 0 | Pa
AS LEFT (PREDICTED) DATA FOR UPSTREAM (HIGH) TRANSDUCER WITH AS FOUND DATA AND AS LEFT PA AND PM | [5 empty] | AS LEFT (PREDICTED) DATA FOR DOWNSTREAM (LOW) TRANSDUCER WITH AS FOUND DATA AND AS LEFT PA AND PM
Reference | Test | Error | As Left Spec(+) | As Left Spec(-) | Pass/Fail | Reference | Test | Error | As Left Spec(+) | As Left Spec(-) | Pass/Fail
kPa | kPa | kPa | kPa | kPa |  | kPa | kPa | kPa | kPa | kPa
19.8531 | 19.8553 | 0.002239 | 0.0168 | -0.0168 | Pass | 19.8531 | 19.8533 | 0.00017476 | 0.0168 | -0.0168 | Pass
99.9873 | 99.9846 | -0.002728 | 0.0168 | -0.0168 | Pass | 99.9873 | 99.9866 | -0.0007208 | 0.0168 | -0.0168 | Pass
200.7614 | 200.7566 | -0.0047925 | 0.0168 | -0.0168 | Pass | 200.7614 | 200.76 | -0.001438 | 0.0168 | -0.0168 | Pass
299.1284 | 299.1257 | -0.002734 | 0.02393 | -0.02393 | Pass | 299.1284 | 299.1263 | -0.0020642 | 0.02393 | -0.02393 | Pass
449.8313 | 449.8302 | -0.0010764 | 0.035987 | -0.035987 | Pass | 449.8313 | 449.8299 | -0.0013917 | 0.035987 | -0.035987 | Pass
600.266 | 600.2688 | 0.0027838 | 0.048021 | -0.048021 | Pass | 600.266 | 600.2675 | 0.0014798 | 0.048021 | -0.048021 | Pass
449.8311 | 449.8322 | 0.0011236 | 0.035986 | -0.035986 | Pass | 449.8311 | 449.8319 | 0.00080831 | 0.035986 | -0.035986 | Pass
299.1281 | 299.1287 | 0.000566 | 0.02393 | -0.02393 | Pass | 299.1281 | 299.1283 | 0.00023577 | 0.02393 | -0.02393 | Pass
200.761 | 200.7606 | -0.0003925 | 0.0168 | -0.0168 | Pass | 200.761 | 200.762 | 0.00096206 | 0.0168 | -0.0168 | Pass
99.987 | 99.9886 | 0.001572 | 0.0168 | -0.0168 | Pass | 99.987 | 99.9876 | 0.0005792 | 0.0168 | -0.0168 | Pass
19.8529 | 19.8563 | 0.003439 | 0.0168 | -0.0168 | Pass | 19.8529 | 19.8543 | 0.0013748 | 0.0168 | -0.0168 | Pass
AS LEFT UPSTREAM (HIGH) TRANSDUCER | [5 empty] | AS LEFT DOWNSTREAM (LOW) TRANSDUCER
Reference | Test | Error | Specification(+) | Specification(-) | Pass/Fail | Reference | Test | Error | Specification(+) | Specification(-) | Pass/Fail
kPa | kPa | kPa | kPa | kPa |  | kPa | kPa | kPa | kPa | kPa
19.8531 | 19.855 | 0.0019 | 0.0168 | -0.0168 | Pass | 19.8531 | 19.854 | 0.0009 | 0.0168 | -0.0168 | Pass
99.9873 | 99.984 | -0.0033 | 0.0168 | -0.0168 | Pass | 99.9873 | 99.986 | -0.0013 | 0.0168 | -0.0168 | Pass
200.7614 | 200.757 | -0.0044 | 0.0168 | -0.0168 | Pass | 200.7614 | 200.76 | -0.0014 | 0.0168 | -0.0168 | Pass
299.1284 | 299.125 | -0.0034 | 0.02393 | -0.02393 | Pass | 299.1284 | 299.127 | -0.0014 | 0.02393 | -0.02393 | Pass
449.8313 | 449.83 | -0.0013 | 0.035987 | -0.035987 | Pass | 449.8313 | 449.83 | -0.0013 | 0.035987 | -0.035987 | Pass
600.266 | 600.269 | 0.003 | 0.048021 | -0.048021 | Pass | 600.266 | 600.267 | 0.001 | 0.048021 | -0.048021 | Pass
449.8311 | 449.832 | 0.0009 | 0.035986 | -0.035986 | Pass | 449.8311 | 449.832 | 0.0009 | 0.035986 | -0.035986 | Pass
299.1281 | 299.129 | 0.0009 | 0.02393 | -0.02393 | Pass | 299.1281 | 299.128 | -0.0001 | 0.02393 | -0.02393 | Pass
200.761 | 200.761 | 0 | 0.0168 | -0.0168 | Pass | 200.761 | 200.762 | 0.001 | 0.0168 | -0.0168 | Pass
99.987 | 99.989 | 0.002 | 0.0168 | -0.0168 | Pass | 99.987 | 99.987 | 0 | 0.0168 | -0.0168 | Pass
19.8529 | 19.857 | 0.0041 | 0.0168 | -0.0168 | Pass | 19.8529 | 19.854 | 0.0011 | 0.0168 | -0.0168 | Pass
As Left Specification: | ±(0.008% of reading or 0.0024% of FS, whichever is greater) |  |  |  |  | As Left Specification: | ±(0.008% of reading or 0.0024% of FS, whichever is greater)
New Pressure Adder (PA): | -6.5709 |  |  |  |  | New Pressure Adder (PA): | 137.5999
New Pressure Multiplier (PM): | 1 |  |  |  |  | New Pressure Multiplier (PM): | 1.0001
AutoZ (zOffset): | 0 | Pa |  |  |  | AutoZ (zOffset): | 0 | Pa