## Existing Format_Changes
Tables
Nitrogen Ratio = 1 | N2 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.5 | 1 | 2 | 2 | 4 | 6.5 | 7 | 12 | 20
2E1-S | 20 | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
5E1-S | 50 | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
1E2-S | 100 | 2 | 5 | 10 | 15 | 20 | 40 | 60 | 70 | 120 | 200
2E2-S | 200 | 4 | 10 | 20 | 28 | 40 | 80 | 120 | 140 | 240 | 400
5E2-S | 500 | 10 | 25 | 50 | 67 | 100 | 200 | 300 | 350 | 600 | 1000
1E3-S | 1000 | 20 | 50 | 100 | 129 | 200 | 400 | 600 | 700 | 1200 | 2000
2E3-S | 2000 | 40 | 100 | 200 | 248 | 400 | 800 | 1200 | 1400 | 2400 | 4000
5E3-S | 5000 | 100 | 250 | 500 | 596 | 1000 | 2000 | 3000 | 3500 | 6000 | 10000
1E4-S | 10000 | 200 | 500 | 1000 | 1173 | 2000 | 4000 | 6000 | 7000 | 12000 | 20000
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum).
Air Ratio = 0.983 | Air | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.5 | 1 | 2 | 2 | 3.95 | 5.9 | 6.9 | 11.8 | 19.7
2E1-S | 20 | 0.4 | 1 | 2 | 3.4 | 3.9 | 7.85 | 11.8 | 13.8 | 23.6 | 39.3
5E1-S | 50 | 1 | 2.5 | 4.9 | 7.6 | 9.8 | 19.65 | 29.5 | 34.4 | 59 | 98.3
1E2-S | 100 | 2 | 4.9 | 9.8 | 15.2 | 19.7 | 39.35 | 59 | 68.8 | 118 | 196.7
2E2-S | 200 | 3.9 | 9.8 | 19.7 | 27.4 | 39.3 | 78.65 | 118 | 137.7 | 236 | 393.4
5E2-S | 500 | 9.8 | 24.6 | 49.2 | 67.1 | 98.3 | 196.675 | 295.05 | 344.2 | 590.1 | 983.5
1E3-S | 1000 | 19.7 | 49.2 | 98.3 | 126.7 | 196.7 | 393.375 | 590.05 | 688.4 | 1180.2 | 1967
2E3-S | 2000 | 39.3 | 98.3 | 196.7 | 243.9 | 393.4 | 786.8 | 1180.2 | 1376.9 | 2360.4 | 3933.9
5E3-S | 5000 | 98.3 | 245.9 | 491.7 | 585.2 | 983.5 | 1966.975 | 2950.45 | 3442.2 | 5900.9 | 9834.8
1E4-S | 10000 | 196.7 | 491.7 | 983.5 | 1170.3 | 1967 | 3933.95 | 5900.9 | 6884.4 | 11801.8 | 19669.7
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum).
Argon Ratio = 0.837 | Ar | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.4 | 0.8 | 1.7 | 1.7 | 3.275 | 4.85 | 5.9 | 10 | 16.7
2E1-S | 20 | 0.3 | 0.8 | 1.7 | 3 | 3.3 | 6.2438 | 9.1875 | 11.7 | 20.1 | 33.5
5E1-S | 50 | 0.8 | 2.1 | 4.2 | 6.5 | 8.4 | 15.7125 | 23.025 | 29.3 | 50.2 | 83.7
1E2-S | 100 | 1.7 | 4.2 | 8.4 | 12.9 | 16.7 | 31.3687 | 46.0375 | 58.6 | 100.5 | 167.4
2E2-S | 200 | 3.3 | 8.4 | 16.7 | 23.3 | 33.5 | 62.7938 | 92.0875 | 117.2 | 200.9 | 334.9
5E2-S | 500 | 8.4 | 20.9 | 41.9 | 57.1 | 83.7 | 156.9563 | 230.2125 | 293 | 502.3 | 837.2
1E3-S | 1000 | 16.7 | 41.9 | 83.7 | 107.8 | 167.4 | 313.9125 | 460.425 | 586 | 1004.6 | 1674.4
2E3-S | 2000 | 33.5 | 83.7 | 167.4 | 207.6 | 334.9 | 627.9188 | 920.9375 | 1172.1 | 2009.3 | 3348.8
5E3-S | 5000 | 83.7 | 209.3 | 418.6 | 498.2 | 837.2 | 1569.75 | 2302.3 | 2930.2 | 5023.2 | 8372
1E4-S | 10000 | 167.4 | 418.6 | 837.2 | 996.2 | 1674.4 | 3139.5 | 4604.6 | 5860.4 | 10046.4 | 16744
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum).
Helium Ratio = 2.647 | He | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.5 | 1.3 | 2.6 | 9.4 | 5.3 | 9.7137 | 14.1273 | 18.5 | 31.8 | 52.9
2E1-S | 20 | 1.1 | 2.6 | 5.3 | 13.1 | 10.6 | 21.2 | 29.15 | 37.1 | 63.5 | 105.9
5E1-S | 50 | 2.6 | 6.6 | 13.2 | 25.7 | 26.5 | 52.95 | 72.775 | 92.6 | 158.8 | 264.7
1E2-S | 100 | 5.3 | 13.2 | 26.5 | 51.4 | 52.9 | 105.85 | 145.575 | 185.3 | 317.6 | 529.4
2E2-S | 200 | 10.6 | 26.5 | 52.9 | 91.5 | 105.9 | 211.775 | 291.1875 | 370.6 | 635.2 | 1058.7
5E2-S | 500 | 26.5 | 66.2 | 132.3 | 199.4 | 264.7 | 529.375 | 727.8875 | 926.4 | 1588.1 | 2646.8
1E3-S | 1000 | 52.9 | 132.3 | 264.7 | 398.7 | 529.4 | 1058.75 | 1455.775 | 1852.8 | 3176.2 | 5293.6
2E3-S | 2000 | 105.9 | 264.7 | 529.4 | 695.1 | 1058.7 | 2117.425 | 2911.4625 | 3705.5 | 6352.3 | 10587.2
5E3-S | 5000 | 264.7 | 661.7 | 1323.4 | 1737.8 | 2646.8 | 5293.6 | 7278.7 | 9263.8 | 15880.9 | 26468.1
1E4-S | 10000 | 529.4 | 1323.4 | 2646.8 | 3281 | 5293.6 | 10587.225 | 14557.4625 | 18527.7 | 31761.7 | 52936.2
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum).
Carbon Dioxide Ratio = 0.795 | CO2 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[4] | n/a[4]
2E1-S | 20 | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
5E1-S | 50 | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
1E2-S | 100 | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
2E2-S | 200 | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
5E2-S | 500 | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.05 | 238.6 | 278.4 | n/a[4] | n/a[4]
1E3-S | 1000 | 15.9 | 39.8 | 79.5 | 102.4 | 159.1 | 318.15 | 477.2 | 556.7 | n/a[4] | n/a[4]
2E3-S | 2000 | 31.8 | 79.5 | 159.1 | 189.3 | 318.1 | 636.225 | 954.35 | 1113.4 | n/a[4] | n/a[4]
5E3-S | 5000 | 79.5 | 198.8 | 397.7 | 473.2 | 795.3 | 1590.625 | 2385.95 | 2783.6 | n/a[4] | n/a[4]
1E4-S | 10000 | 159.1 | 397.7 | 795.3 | 914.1 | 1590.6 | 3181.25 | 4771.9 | 5567.2 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Sulfur Hexafluoride Ratio = 0.435 | SF6 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.75 | 2.6 | 3 | n/a[4] | n/a[4]
2E1-S | 20 | 0.2 | 0.4 | 0.9 | 1.4 | 1.7 | 3.45 | 5.2 | 6.1 | n/a[4] | n/a[4]
5E1-S | 50 | 0.4 | 1.1 | 2.2 | 3.1 | 4.3 | 8.675 | 13.05 | 15.2 | n/a[4] | n/a[4]
1E2-S | 100 | 0.9 | 2.2 | 4.3 | 5.9 | 8.7 | 17.375 | 26.05 | 30.4 | n/a[4] | n/a[4]
2E2-S | 200 | 1.7 | 4.3 | 8.7 | 11.4 | 17.4 | 34.8 | 52.2 | 60.9 | n/a[4] | n/a[4]
5E2-S | 500 | 4.3 | 10.9 | 21.7 | 26.9 | 43.5 | 86.975 | 130.45 | 152.2 | n/a[4] | n/a[4]
1E3-S | 1000 | 8.7 | 21.7 | 43.5 | 53.8 | 87 | 173.925 | 260.85 | 304.3 | n/a[4] | n/a[4]
2E3-S | 2000 | 17.4 | 43.5 | 87 | 100 | 173.9 | 347.825 | 521.75 | 608.7 | n/a[4] | n/a[4]
5E3-S | 5000 | 43.5 | 108.7 | 217.4 | 249.9 | 434.8 | 869.55 | 1304.3 | 1521.7 | n/a[4] | n/a[4]
1E4-S | 10000 | 87 | 217.4 | 434.8 | 499.7 | 869.5 | 1739.05 | 2608.6 | 3043.4 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Xenon Ratio = 0.460 | Xe | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.1 | 0.2 | 0.4 | 0.8 | 0.7 | 1.6 | 2.5 | 3.2 | n/a[4] | n/a[4]
2E1-S | 20 | 0.1 | 0.4 | 0.7 | 1.4 | 1.5 | 3.25 | 5 | 6.4 | n/a[4] | n/a[4]
5E1-S | 50 | 0.4 | 0.9 | 1.8 | 3.6 | 3.6 | 8.1 | 12.6 | 16.1 | n/a[4] | n/a[4]
1E2-S | 100 | 0.7 | 1.8 | 3.6 | 6.5 | 7.3 | 16.25 | 25.2 | 32.2 | n/a[4] | n/a[4]
2E2-S | 200 | 1.5 | 3.6 | 7.3 | 12.9 | 14.6 | 32.525 | 50.45 | 64.5 | n/a[4] | n/a[4]
5E2-S | 500 | 3.6 | 9.1 | 18.2 | 29.7 | 36.4 | 81.25 | 126.1 | 161.2 | n/a[4] | n/a[4]
1E3-S | 1000 | 7.3 | 18.2 | 36.4 | 59.3 | 72.8 | 162.475 | 252.15 | 322.3 | n/a[4] | n/a[4]
2E3-S | 2000 | 14.6 | 36.4 | 72.8 | 109.6 | 145.6 | 324.95 | 504.3 | 644.6 | n/a[4] | n/a[4]
5E3-S | 5000 | 36.4 | 91 | 182 | 267.2 | 364 | 812.375 | 1260.75 | 1611.5 | n/a[4] | n/a[4]
1E4-S | 10000 | 72.8 | 182 | 364 | 529.2 | 728 | 1624.75 | 2521.5 | 3223.1 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Butane Ratio = 0.680 | C4H10 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.1 | 0.3 | 0.7 | 1 | 1.4 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
2E1-S | 20 | 0.3 | 0.7 | 1.4 | 1.9 | 2.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
5E1-S | 50 | 0.7 | 1.7 | 3.4 | 4.8 | 6.8 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
1E2-S | 100 | 1.4 | 3.4 | 6.8 | 9 | 13.6 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
2E2-S | 200 | 2.7 | 6.8 | 13.6 | 18 | 27.2 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
5E2-S | 500 | 6.8 | 17 | 34 | 42.2 | 68 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
1E3-S | 1000 | 13.6 | 34 | 68 | 84.4 | 136 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
2E3-S | 2000 | 27.2 | 68 | 136 | 157 | 272 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
5E3-S | 5000 | 68 | 170 | 340 | 392.4 | 679.9 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
1E4-S | 10000 | 136 | 340 | 679.9 | 784.9 | 1359.8 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Ethane Ratio = 0.960 | C2H6 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.5 | 1 | 1.5 | 1.9 | 3.825 | 5.75 | 6.7 | n/a[4] | n/a[4]
2E1-S | 20 | 0.4 | 1 | 1.9 | 3 | 3.8 | 7.65 | 11.5 | 13.4 | n/a[4] | n/a[4]
5E1-S | 50 | 1 | 2.4 | 4.8 | 6.7 | 9.6 | 19.2 | 28.8 | 33.6 | n/a[4] | n/a[4]
1E2-S | 100 | 1.9 | 4.8 | 9.6 | 13.4 | 19.2 | 38.4 | 57.6 | 67.2 | n/a[4] | n/a[4]
2E2-S | 200 | 3.8 | 9.6 | 19.2 | 25.2 | 38.4 | 76.825 | 115.25 | 134.5 | n/a[4] | n/a[4]
5E2-S | 500 | 9.6 | 24 | 48 | 61.9 | 96 | 192.05 | 288.1 | 336.1 | n/a[4] | n/a[4]
1E3-S | 1000 | 19.2 | 48 | 96 | 118.7 | 192.1 | 384.175 | 576.25 | 672.3 | n/a[4] | n/a[4]
2E3-S | 2000 | 38.4 | 96 | 192.1 | 228.6 | 384.2 | 768.35 | 1152.5 | 1344.6 | n/a[4] | n/a[4]
5E3-S | 5000 | 96 | 240.1 | 480.2 | 551.9 | 960.4 | 1920.8 | 2881.2 | 3361.4 | n/a[4] | n/a[4]
1E4-S | 10000 | 192.1 | 480.2 | 960.4 | 1103.8 | 1920.8 | 3841.6 | 5762.4 | 6722.8 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Ethylene Ratio = 0.996 | C2H4 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.5 | 1 | 1.7 | 2 | 4 | 6 | 7 | n/a[4] | n/a[4]
2E1-S | 20 | 0.4 | 1 | 2 | 3.1 | 4 | 7.975 | 11.95 | 13.9 | n/a[4] | n/a[4]
5E1-S | 50 | 1 | 2.5 | 5 | 7.5 | 10 | 19.95 | 29.9 | 34.9 | n/a[4] | n/a[4]
1E2-S | 100 | 2 | 5 | 10 | 13.9 | 19.9 | 39.825 | 59.75 | 69.7 | n/a[4] | n/a[4]
2E2-S | 200 | 4 | 10 | 19.9 | 27.7 | 39.8 | 79.65 | 119.5 | 139.4 | n/a[4] | n/a[4]
5E2-S | 500 | 10 | 24.9 | 49.8 | 64.2 | 99.6 | 199.15 | 298.7 | 348.5 | n/a[4] | n/a[4]
1E3-S | 1000 | 19.9 | 49.8 | 99.6 | 128.2 | 199.2 | 398.325 | 597.45 | 697 | n/a[4] | n/a[4]
2E3-S | 2000 | 39.8 | 99.6 | 199.2 | 237 | 398.3 | 796.625 | 1194.95 | 1394.1 | n/a[4] | n/a[4]
5E3-S | 5000 | 99.6 | 248.9 | 497.9 | 572.2 | 995.8 | 1991.575 | 2987.35 | 3485.2 | n/a[4] | n/a[4]
1E4-S | 10000 | 199.2 | 497.9 | 995.8 | 1144.4 | 1991.6 | 3983.15 | 5974.7 | 6970.5 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Hydrogen Ratio = 3.730 | H2 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.7 | 1.9 | 3.7 | 8.3 | 7.5 | 14.925 | 22.35 | 26.1 | 44.8 | 74.6
2E1-S | 20 | 1.5 | 3.7 | 7.5 | 14.5 | 14.9 | 29.825 | 44.75 | 52.2 | 89.5 | 149.2
5E1-S | 50 | 3.7 | 9.3 | 18.6 | 36.2 | 37.3 | 74.575 | 111.85 | 130.5 | 223.8 | 373
1E2-S | 100 | 7.5 | 18.6 | 37.3 | 62.5 | 74.6 | 149.2 | 223.8 | 261.1 | 447.6 | 746
2E2-S | 200 | 14.9 | 37.3 | 74.6 | 114.5 | 149.2 | 298.4 | 447.6 | 522.2 | 895.2 | 1491.9
5E2-S | 500 | 37.3 | 93.2 | 186.5 | 280.9 | 373 | 745.95 | 1118.9 | 1305.4 | 2237.9 | 3729.8
1E3-S | 1000 | 74.6 | 186.5 | 373 | 508.7 | 746 | 1491.95 | 2237.9 | 2610.9 | 4475.8 | 7459.6
2E3-S | 2000 | 149.2 | 373 | 746 | 979.6 | 1491.9 | 2983.825 | 4475.75 | 5221.7 | 8951.5 | 14919.2
5E3-S | 5000 | 373 | 932.4 | 1864.9 | 2311.7 | 3729.8 | 7459.6 | 11189.4 | 13054.3 | 22378.8 | 37297.9
1E4-S | 10000 | 746 | 1864.9 | 3729.8 | 4623.4 | 7459.6 | 14919.2 | 22378.8 | 26108.6 | 44757.5 | 74595.9
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum).
Methane Ratio = 1.320 | CH4 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.3 | 0.7 | 1.3 | 2.6 | 2.6 | 5.25 | 7.9 | 9.2 | 15.8 | 26.4
2E1-S | 20 | 0.5 | 1.3 | 2.6 | 4.4 | 5.3 | 10.575 | 15.85 | 18.5 | 31.7 | 52.8
5E1-S | 50 | 1.3 | 3.3 | 6.6 | 10.2 | 13.2 | 26.4 | 39.6 | 46.2 | 79.2 | 132
1E2-S | 100 | 2.6 | 6.6 | 13.2 | 20.1 | 26.4 | 52.8 | 79.2 | 92.4 | 158.4 | 264
2E2-S | 200 | 5.3 | 13.2 | 26.4 | 36.7 | 52.8 | 105.6 | 158.4 | 184.8 | 316.8 | 528.1
5E2-S | 500 | 13.2 | 33 | 66 | 88.2 | 132 | 264.025 | 396.05 | 462.1 | 792.1 | 1320.2
1E3-S | 1000 | 26.4 | 66 | 132 | 170 | 264 | 528.05 | 792.1 | 924.1 | 1584.2 | 2640.3
2E3-S | 2000 | 52.8 | 132 | 264 | 327.3 | 528.1 | 1056.15 | 1584.2 | 1848.2 | 3168.4 | 5280.6
5E3-S | 5000 | 132 | 330 | 660.1 | 785.5 | 1320.2 | 2640.325 | 3960.45 | 4620.5 | 7920.9 | 13201.5
1E4-S | 10000 | 264 | 660.1 | 1320.2 | 1517.2 | 2640.3 | 5280.625 | 7920.95 | 9241.1 | 15841.8 | 26403
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum).
Propane Ratio = 0.789 | C3H8 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.4 | 0.8 | 1.3 | 1.6 | 3.15 | 4.7 | 5.5 | n/a[4] | n/a[4]
2E1-S | 20 | 0.3 | 0.8 | 1.6 | 2.3 | 3.2 | 6.325 | 9.45 | 11 | n/a[4] | n/a[4]
5E1-S | 50 | 0.8 | 2 | 3.9 | 5.5 | 7.9 | 15.775 | 23.65 | 27.6 | n/a[4] | n/a[4]
1E2-S | 100 | 1.6 | 3.9 | 7.9 | 10.5 | 15.8 | 31.55 | 47.3 | 55.2 | n/a[4] | n/a[4]
2E2-S | 200 | 3.2 | 7.9 | 15.8 | 20.8 | 31.5 | 63.075 | 94.65 | 110.4 | n/a[4] | n/a[4]
5E2-S | 500 | 7.9 | 19.7 | 39.4 | 48.8 | 78.9 | 157.775 | 236.65 | 276.1 | n/a[4] | n/a[4]
1E3-S | 1000 | 15.8 | 39.4 | 78.9 | 97.5 | 157.7 | 315.475 | 473.25 | 552.1 | n/a[4] | n/a[4]
2E3-S | 2000 | 31.5 | 78.9 | 157.7 | 181.4 | 315.5 | 630.975 | 946.45 | 1104.2 | n/a[4] | n/a[4]
5E3-S | 5000 | 78.9 | 197.2 | 394.4 | 453.2 | 788.7 | 1577.425 | 2366.15 | 2760.5 | n/a[4] | n/a[4]
1E4-S | 10000 | 157.7 | 394.4 | 788.7 | 906.5 | 1577.4 | 3154.85 | 4732.3 | 5521 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Carbon Tetrafluoride Ratio = 0.563 | CF4 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.1 | 0.3 | 0.6 | 0.9 | 1.1 | 2.225 | 3.35 | 3.9 | n/a[4] | n/a[4]
2E1-S | 20 | 0.2 | 0.6 | 1.1 | 1.8 | 2.3 | 4.525 | 6.75 | 7.9 | n/a[4] | n/a[4]
5E1-S | 50 | 0.6 | 1.4 | 2.8 | 4.1 | 5.6 | 11.25 | 16.9 | 19.7 | n/a[4] | n/a[4]
1E2-S | 100 | 1.1 | 2.8 | 5.6 | 7.9 | 11.3 | 22.525 | 33.75 | 39.4 | n/a[4] | n/a[4]
2E2-S | 200 | 2.3 | 5.6 | 11.3 | 15.7 | 22.5 | 45.025 | 67.55 | 78.8 | n/a[4] | n/a[4]
5E2-S | 500 | 5.6 | 14.1 | 28.1 | 36.3 | 56.3 | 112.575 | 168.85 | 197 | n/a[4] | n/a[4]
1E3-S | 1000 | 11.3 | 28.1 | 56.3 | 69.9 | 112.6 | 225.15 | 337.7 | 394 | n/a[4] | n/a[4]
2E3-S | 2000 | 22.5 | 56.3 | 112.6 | 134.1 | 225.1 | 450.275 | 675.45 | 788 | n/a[4] | n/a[4]
5E3-S | 5000 | 56.3 | 140.7 | 281.4 | 323.4 | 562.9 | 1125.75 | 1688.6 | 1970 | n/a[4] | n/a[4]
1E4-S | 10000 | 112.6 | 281.4 | 562.9 | 646.9 | 1125.7 | 2251.45 | 3377.2 | 3940.1 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Hexafluoroethane Ratio = 0.447 | C2F6 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.775 | 2.65 | 3.1 | n/a[4] | n/a[4]
2E1-S | 20 | 0.2 | 0.4 | 0.9 | 1.3 | 1.8 | 3.6 | 5.4 | 6.3 | n/a[4] | n/a[4]
5E1-S | 50 | 0.4 | 1.1 | 2.2 | 3.2 | 4.5 | 8.975 | 13.45 | 15.7 | n/a[4] | n/a[4]
1E2-S | 100 | 0.9 | 2.2 | 4.5 | 5.9 | 8.9 | 17.875 | 26.85 | 31.3 | n/a[4] | n/a[4]
2E2-S | 200 | 1.8 | 4.5 | 8.9 | 11.8 | 17.9 | 35.775 | 53.65 | 62.6 | n/a[4] | n/a[4]
5E2-S | 500 | 4.5 | 11.2 | 22.4 | 27.6 | 44.7 | 89.45 | 134.2 | 156.6 | n/a[4] | n/a[4]
1E3-S | 1000 | 8.9 | 22.4 | 44.7 | 55.3 | 89.5 | 178.95 | 268.4 | 313.1 | n/a[4] | n/a[4]
2E3-S | 2000 | 17.9 | 44.7 | 89.5 | 102.9 | 178.9 | 357.85 | 536.8 | 626.3 | n/a[4] | n/a[4]
5E3-S | 5000 | 44.7 | 111.8 | 223.7 | 257.1 | 447.3 | 894.65 | 1342 | 1565.7 | n/a[4] | n/a[4]
1E4-S | 10000 | 89.5 | 223.7 | 447.3 | 514.1 | 894.7 | 1789.35 | 2684 | 3131.3 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Trifluoromethane Ratio = 0.629 | CHF3 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.1 | 0.3 | 0.6 | 1 | 1.3 | 2.525 | 3.75 | 4.4 | n/a[4] | n/a[4]
2E1-S | 20 | 0.3 | 0.6 | 1.3 | 2 | 2.5 | 5.025 | 7.55 | 8.8 | n/a[4] | n/a[4]
5E1-S | 50 | 0.6 | 1.6 | 3.1 | 4.4 | 6.3 | 12.575 | 18.85 | 22 | n/a[4] | n/a[4]
1E2-S | 100 | 1.3 | 3.1 | 6.3 | 8.8 | 12.6 | 25.2 | 37.8 | 44.1 | n/a[4] | n/a[4]
2E2-S | 200 | 2.5 | 6.3 | 12.6 | 17.2 | 25.2 | 50.35 | 75.5 | 88.1 | n/a[4] | n/a[4]
5E2-S | 500 | 6.3 | 15.7 | 31.5 | 40.6 | 62.9 | 125.85 | 188.8 | 220.3 | n/a[4] | n/a[4]
1E3-S | 1000 | 12.6 | 31.5 | 62.9 | 77.9 | 125.9 | 251.775 | 377.65 | 440.6 | n/a[4] | n/a[4]
2E3-S | 2000 | 25.2 | 62.9 | 125.9 | 149.9 | 251.8 | 503.525 | 755.25 | 881.1 | n/a[4] | n/a[4]
5E3-S | 5000 | 62.9 | 157.3 | 314.7 | 361.7 | 629.4 | 1258.75 | 1888.1 | 2202.8 | n/a[4] | n/a[4]
1E4-S | 10000 | 125.9 | 314.7 | 629.4 | 723.3 | 1258.8 | 2517.525 | 3776.25 | 4405.6 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Carbon Monoxide Ratio = 1.000 | CO | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.5 | 1 | 2 | 2 | 4 | 6 | 7 | 12 | 20
2E1-S | 20 | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
5E1-S | 50 | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
1E2-S | 100 | 2 | 5 | 10 | 15.4 | 20 | 40 | 60 | 70 | 120 | 200
2E2-S | 200 | 4 | 10 | 20 | 27.8 | 40 | 80 | 120 | 140 | 240 | 400
5E2-S | 500 | 10 | 25 | 50 | 68.3 | 100 | 200 | 300 | 350 | 600 | 1000
1E3-S | 1000 | 20 | 50 | 100 | 128.7 | 200 | 400 | 600 | 700 | 1200 | 1999.9
2E3-S | 2000 | 40 | 100 | 200 | 247.9 | 400 | 799.975 | 1199.95 | 1399.9 | 2399.9 | 3999.9
5E3-S | 5000 | 100 | 250 | 500 | 595 | 1000 | 1999.95 | 2999.9 | 3499.9 | 5999.8 | 9999.6
1E4-S | 10000 | 200 | 500 | 1000 | 1190 | 1999.9 | 3999.825 | 5999.75 | 6999.7 | 11999.6 | 19999.3
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum).
Nitrous Oxide Ratio = 0.795 | N2O | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[4] | n/a[4]
2E1-S | 20 | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
5E1-S | 50 | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
1E2-S | 100 | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
2E2-S | 200 | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
5E2-S | 500 | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.025 | 238.55 | 278.3 | n/a[4] | n/a[4]
1E3-S | 1000 | 15.9 | 39.8 | 79.5 | 102.4 | 159 | 318.05 | 477.1 | 556.6 | n/a[4] | n/a[4]
2E3-S | 2000 | 31.8 | 79.5 | 159 | 189.3 | 318 | 636.075 | 954.15 | 1113.2 | n/a[4] | n/a[4]
5E3-S | 5000 | 79.5 | 198.8 | 397.6 | 473.1 | 795.1 | 1590.225 | 2385.35 | 2782.9 | n/a[4] | n/a[4]
1E4-S | 10000 | 159 | 397.6 | 795.1 | 913.7 | 1590.2 | 3180.45 | 4770.7 | 5565.8 | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Octafluorocyclobutane1 Ratio = 0.367 | C4F8 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.1 | 0.2 | 0.4 | 0.6 | 0.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
2E1-S | 20 | 0.1 | 0.4 | 0.7 | 1.1 | 1.5 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
5E1-S | 50 | 0.4 | 0.9 | 1.8 | 2.4 | 3.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
1E2-S | 100 | 0.7 | 1.8 | 3.7 | 4.8 | 7.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
2E2-S | 200 | 1.5 | 3.7 | 7.3 | 9.2 | 14.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
5E2-S | 500 | 3.7 | 9.2 | 18.3 | 22.7 | 36.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
1E3-S | 1000 | 7.3 | 18.3 | 36.7 | 43.8 | 73.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
2E3-S | 2000 | 14.7 | 36.7 | 73.3 | 84.3 | 146.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
5E3-S | 5000 | 36.7 | 91.7 | 183.3 | 210.7 | 366.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
1E4-S | 10000 | 73.3 | 183.3 | 366.7 | 421.4 | 733.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Oxygen Ratio = 0.935 | O2 | molbloc-S Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2]
Designator | KF [sccm/kPa] | 20 kPa (3 psia) | 50 kPa (7 psia) | 100 kPa (15 psia) | Minimum without vacuum[3] | 200 kPa (30 psia) | 400 kPa (58 psia) | 600 kPa (87 psia) | 700 kPa (100 psia) | 1.20 MPa (174 psia) | 2 MPa (290 psia)
1E1-S | 10 | 0.2 | 0.5 | 0.9 | 1.9 | 1.9 | 3.75 | 5.6 | 6.5 | 11.2 | 18.7
2E1-S | 20 | 0.4 | 0.9 | 1.9 | 3.4 | 3.7 | 7.475 | 11.25 | 13.1 | 22.5 | 37.4
5E1-S | 50 | 0.9 | 2.3 | 4.7 | 7.3 | 9.4 | 18.725 | 28.05 | 32.7 | 56.1 | 93.5
1E2-S | 100 | 1.9 | 4.7 | 9.4 | 14.4 | 18.7 | 37.425 | 56.15 | 65.5 | 112.3 | 187.1
2E2-S | 200 | 3.7 | 9.4 | 18.7 | 26.4 | 37.4 | 74.825 | 112.25 | 131 | 224.5 | 374.2
5E2-S | 500 | 9.4 | 23.4 | 46.8 | 63.8 | 93.5 | 187.075 | 280.65 | 327.4 | 561.3 | 935.4
1E3-S | 1000 | 18.7 | 46.8 | 93.5 | 120.4 | 187.1 | 374.175 | 561.25 | 654.8 | 1122.5 | 1870.8
2E3-S | 2000 | 37.4 | 93.5 | 187.1 | 231.9 | 374.2 | 748.35 | 1122.5 | 1309.6 | 2245 | 3741.7
5E3-S | 5000 | 93.5 | 233.9 | 467.7 | 556.6 | 935.4 | 1870.825 | 2806.25 | 3274 | 5612.5 | 9354.2
1E4-S | 10000 | 187.1 | 467.7 | 935.4 | 1113.1 | 1870.8 | 3741.65 | 5612.5 | 6547.9 | 11225 | 18708.3
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum).

## New_Min Max_5 gases
molbloc-S Pressure Dependent Flow Rates: Min/Max Mass Flow Rate (slm @ 0C) (Kf designator = sccm per kPa):[1][2][10]
|  |  |  | 1E1-S |  | 2E1-S |  | 5E1-S |  | 1E2-S |  | 2E2-S |  | 5E2-S |  | 1E3-S |  | 2E3-S |  | 5E3-S |  | 1E4-S
|  |  |  | 1E1-S Kf = | 10 | 1E1-S Kf = | 20 | 1E1-S Kf = | 50 | 1E1-S Kf = | 100 | 1E1-S Kf = | 200 | 1E1-S Kf = | 500 | 1E1-S Kf = | 1000 | 1E1-S Kf = | 2000 | 1E1-S Kf = | 5000 | 1E1-S Kf = | 10000
|  | Pressure Calibration type |  | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max
N2 |  | Low Pressure (LP) | (with vacuum)[4] | 0.2 | 2 | 0.4 | 4 | 1 | 10 | 2 | 20 | 4 | 40 | 10 | 100 | 20 | 200 | 40 | 400 | 100 | 1000 | 200 | 2000
Ratio = | 1 |  | (no vacuum)[3][5] | 2 | 2 | 3.5 | 4 | 7.7 | 10 | 15 | 20 | 28 | 40 | 67 | 100 | 129 | 200 | 248 | 400 | 596 | 1000 | 1173 | 2000
|  | Standard Pressure (SP) | (with vacuum)[6] | 0.5 | 6 | 1 | 12 | 2.5 | 30 | 5 | 60 | 10 | 120 | 25 | 300 | 50 | 600 | 100 | 1200 | 250 | 3000 | 500 | 6000
|  |  | (no vacuum)[3][7] | 2 | 6 | 3.5 | 12 | 7.7 | 30 | 15 | 60 | 28 | 120 | 67 | 300 | 129 | 600 | 248 | 1200 | 596 | 3000 | 1173 | 6000
|  | High Pressure (HP)[8][9] |  | 2 | 20 | 4 | 40 | 10 | 100 | 20 | 200 | 40 | 400 | 100 | 1000 | 200 | 2000 | 400 | 4000 | 1000 | 10000 | 2000 | 20000
Air |  | Low Pressure (LP) | (with vacuum)[4] | 0.1966 | 1.966 | 0.3932 | 3.932 | 0.983 | 9.83 | 1.966 | 19.66 | 3.932 | 39.32 | 9.83 | 98.3 | 19.66 | 196.6 | 39.32 | 393.2 | 98.3 | 983 | 196.6 | 1966
Ratio = | 0.983 |  | (no vacuum)[3][5] | 2 | 1.966 | 3.4 | 3.932 | 7.6 | 9.83 | 15.2 | 19.66 | 27.4 | 39.32 | 67.1 | 98.3 | 126.7 | 196.6 | 243.9 | 393.2 | 585.2 | 983 | 1170.3 | 1966
|  | Standard Pressure (SP) | (with vacuum)[6] | 0.4915 | 5.898 | 0.983 | 11.796 | 2.4575 | 29.49 | 4.915 | 58.98 | 9.83 | 117.96 | 24.575 | 294.9 | 49.15 | 589.8 | 98.3 | 1179.6 | 245.75 | 2949 | 491.5 | 5898
|  |  | (no vacuum)[3][7] | 2 | 5.898 | 3.4 | 11.796 | 7.6 | 29.49 | 15.2 | 58.98 | 27.4 | 117.96 | 67.1 | 294.9 | 126.7 | 589.8 | 243.9 | 1179.6 | 585.2 | 2949 | 1170.3 | 5898
|  | High Pressure (HP)[8][9] |  | 1.966 | 19.66 | 3.932 | 39.32 | 9.83 | 98.3 | 19.66 | 196.6 | 39.32 | 393.2 | 98.3 | 983 | 196.6 | 1966 | 393.2 | 3932 | 983 | 9830 | 1966 | 19660
Argon |  | Low Pressure (LP) | (with vacuum)[4] | 0.1674 | 1.674 | 0.3348 | 3.348 | 0.837 | 8.37 | 1.674 | 16.74 | 3.348 | 33.48 | 8.37 | 83.7 | 16.74 | 167.4 | 33.48 | 334.8 | 83.7 | 837 | 167.4 | 1674
Ratio = | 0.837 |  | (no vacuum)[3][5] | 1.7 | 1.674 | 3 | 3.348 | 6.5 | 8.37 | 12.9 | 16.74 | 23.3 | 33.48 | 57.1 | 83.7 | 107.8 | 167.4 | 207.6 | 334.8 | 498.2 | 837 | 996.2 | 1674
|  | Standard Pressure (SP) | (with vacuum)[6] | 0.4185 | 5.022 | 0.837 | 10.044 | 2.0925 | 25.11 | 4.185 | 50.22 | 8.37 | 100.44 | 20.925 | 251.1 | 41.85 | 502.2 | 83.7 | 1004.4 | 209.25 | 2511 | 418.5 | 5022
|  |  | (no vacuum)[3][7] | 1.7 | 5.022 | 3 | 10.044 | 6.5 | 25.11 | 12.9 | 50.22 | 23.3 | 100.44 | 57.1 | 251.1 | 107.8 | 502.2 | 207.6 | 1004.4 | 498.2 | 2511 | 996.2 | 5022
|  | High Pressure (HP)[8][9] |  | 1.674 | 16.74 | 3.348 | 33.48 | 8.37 | 83.7 | 16.74 | 167.4 | 33.48 | 334.8 | 83.7 | 837 | 167.4 | 1674 | 334.8 | 3348 | 837 | 8370 | 1674 | 16740
Helium |  | Low Pressure (LP) | (with vacuum)[4] | 0.5294 | 5.294 | 1.0588 | 10.588 | 2.647 | 26.47 | 5.294 | 52.94 | 10.588 | 105.88 | 26.47 | 264.7 | 52.94 | 529.4 | 105.88 | 1058.8 | 264.7 | 2647 | 529.4 | 5294
Ratio = | 2.647 |  | (no vacuum)[3][5] | NA[11] | NA[11] | NA[11] | NA[11] | 25.7 | 26.47 | 51.4 | 52.94 | 91.5 | 105.88 | 199.4 | 264.7 | 398.7 | 529.4 | 695.1 | 1058.8 | 1737.8 | 2647 | 3281 | 5294
|  | Standard Pressure (SP) | (with vacuum)[6] | 1.3235 | 15.882 | 2.647 | 31.764 | 6.6175 | 79.41 | 13.235 | 158.82 | 26.47 | 317.64 | 66.175 | 794.1 | 132.35 | 1588.2 | 264.7 | 3176.4 | 661.75 | 7941 | 1323.5 | 15882
|  |  | (no vacuum)[3][7] | NA[11] | NA[11] | NA[11] | NA[11] | 25.7 | 79.41 | 51.4 | 158.82 | 91.5 | 317.64 | 199.4 | 794.1 | 398.7 | 1588.2 | 695.1 | 3176.4 | 1737.8 | 7941 | 3281 | 15882
|  | High Pressure (HP)[8][9] |  | 5.294 | 52.94 | 10.588 | 105.88 | 26.47 | 264.7 | 52.94 | 529.4 | 105.88 | 1058.8 | 264.7 | 2647 | 529.4 | 5294 | 1058.8 | 10588 | 2647 | 26470 | 5294 | 52940
CO2 |  | Low Pressure (LP) | (with vacuum)[4] | 0.159 | 1.59 | 0.318 | 3.18 | 0.795 | 7.95 | 1.59 | 15.9 | 3.18 | 31.8 | 7.95 | 79.5 | 15.9 | 159 | 31.8 | 318 | 79.5 | 795 | 159 | 1590
Ratio = | 0.795 |  | (no vacuum)[3][5] | 1.4 | 1.59 | 2.5 | 3.18 | 6.2 | 7.95 | 11.1 | 15.9 | 22.1 | 31.8 | 51.2 | 79.5 | 102.4 | 159 | 189.3 | 318 | 473.2 | 795 | 914.1 | 1590
|  | Standard Pressure (SP) | (with vacuum)[6] | 0.3975 | 4.77 | 0.795 | 9.54 | 1.9875 | 23.85 | 3.975 | 47.7 | 7.95 | 95.4 | 19.875 | 238.5 | 39.75 | 477 | 79.5 | 954 | 198.75 | 2385 | 397.5 | 4770
|  |  | (no vacuum)[3][7] | 1.4 | 4.77 | 2.5 | 9.54 | 6.2 | 23.85 | 11.1 | 47.7 | 22.1 | 95.4 | 51.2 | 238.5 | 102.4 | 477 | 189.3 | 954 | 473.2 | 2385 | 914.1 | 4770
|  | High Pressure (HP)[8][9] |  | 1.59 | 15.9 | 3.18 | 31.8 | 7.95 | 79.5 | 15.9 | 159 | 31.8 | 318 | 79.5 | 795 | 159 | 1590 | 318 | 3180 | 795 | 7950 | 1590 | 15900
[1] Flow values in table are valid only when critical flow is established. Assumes proper pressurized gas supply and/or vacuum is used when needed to achieve the critical flow and that those sources can maintain the given flow rate. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Upstream pressure applied to achieve minimal critical flow rate with atmospheric pressure (approximately 100 kPaa) downstream of molbloc-S (no vacuum applied to the outlet). [4] Low Pressure calibration (LP) (with vacuum) upstream minimum pressure = 20kPaa (3 PSIa) and maximum pressure = 200kPaa (30 PSIa). [5] Low Pressure calibration (LP) (no vacuum) upstream minimum pressure = atmospheric pressure on the outlet and maximum pressure = 200kPaa (30 PSIa). [6] Standard Pressure (SP) (with vacuum) upstream minimum pressure = 50kPaa (7 PSIa) and maximum pressure = 600kPaa (87 PSIa). [7] Standard Pressure (SP) (no vacuum) upstream minimum pressure = atmospheric pressure on the outlet and maximum pressure = 600kPaa (87 PSIa). [8] High Pressure (HP) upstream minimum pressure = 200kPaa (30 PSIa) and maximum pressure = 2MPaa (290 PSIa). [9] Requires Molbox1+S A2M range to reach maximum flow. Molbox1+S A1.4M will have reduced maximum flow. [10] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [11] Requires vacuum to achieve critical flow.
The table above shows gases that were supported for calibration when this document was released. The flow rates for other gases or at different pressures can be calculated using the formula Kf × Ratio × (Pressure in kPaa/1000). Below is a table of known ratio for various gases to help in these calculations. For example, in the above table you will see the 1E1-S has a Kf = 10 and N2 has a Ratio = 1.00 . The minimum pressure for Low Pressure (LP) (with vacuum) is 20 kPaa per the corresponding footnote. So, the formula is 1.00 x 10 x (20/1000) which would equal the 0.2 value shown in the upper left most table values above:
[6 empty] | Gas |  |  |  |  | Ratio
[6 empty] | Nitrogen |  |  |  | N2 | 1
[6 empty] | Air |  |  |  | Air | 0.983
[6 empty] | Argon |  |  |  | Ar | 0.837
[6 empty] | Helium |  |  |  | He | 2.647
[6 empty] | Carbon Dioxide |  |  |  | CO2 | 0.795
[6 empty] | Sulfur Hexafluoride |  |  |  | SF6 | 0.435
[6 empty] | Xenon |  |  |  | Xe | 0.46
[6 empty] | Butane |  |  |  | C4H10 | 0.68
[6 empty] | Ethane |  |  |  | C2H6 | 0.96
[6 empty] | Ethylene |  |  |  | C2H4 | 0.996
[6 empty] | Hydrogen |  |  |  | H2 | 3.73
[6 empty] | Methane |  |  |  | CH4 | 1.32
[6 empty] | Propane |  |  |  | C3H8 | 0.789
[6 empty] | Carbon Tetrafluoride |  |  |  | CF4 | 0.563
[6 empty] | Hexafluoroethane |  |  |  | C2F6 | 0.447
[6 empty] | Trifluoromethane |  |  |  | CHF3 | 0.629
[6 empty] | Carbon Monoxide |  |  |  | CO | 1
[6 empty] | Nitrous Oxide |  |  |  | N2O | 0.795
[6 empty] | Octafluorocyclobutane1 |  |  |  | C4F8 | 0.367
[6 empty] | Oxygen |  |  |  | O2 | 0.935

## Legacy_extended_all gases
The below tables can be used as an example to show the different flow rates that can be achieved by the same molbloc-S in the same gas, but with different Pressure Dependent Calibration Types. Use this example as a reference to help determine flow rates of different molbloc's and also if vacuum will be used or not in the gas tables on following pages. For example, a 5E2-S in N2 with a Standard Pressure calibration allows about 25 to 300 slm@0C but if not going to be using vacuum to allow achieving the lower flow points its useable range will then be limited at about 67 to 300 slm@0C.
Nitrogen Ratio = 1 | N2 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.5 | 1 | 2 | 2
| Low Pressure (LP) (no vacuum) |  |  |  | 2 | 2
| Standard Pressure (SP) (with vacuum) |  | 0.5 | 1 | 2 | 2 | 4 | 6.5
| Standard Pressure (SP) (no vacuum) |  |  |  | 2 | 2 | 4 | 6.5
| High Pressure (HP) |  |  |  | 2 | 2 | 4 | 6.5 | 7 | 12 | 20
2E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
| Low Pressure (LP) (no vacuum) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
| Standard Pressure (SP) (with vacuum) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
| Standard Pressure (SP) (no vacuum) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
| High Pressure (HP) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
5E1-S | Low Pressure (LP) (with vacuum) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
| Low Pressure (LP) (no vacuum) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
| Standard Pressure (SP) (with vacuum) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
| Standard Pressure (SP) (no vacuum) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
| High Pressure (HP) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
1E2-S | Low Pressure (LP) (with vacuum) | 2 | 5 | 10 | 15 | 20 | 40 | 60 | 70 | 120 | 200
| Low Pressure (LP) (no vacuum) | 2 | 5 | 10 | 15 | 20 | 40 | 60 | 70 | 120 | 200
| Standard Pressure (SP) (with vacuum) | 2 | 5 | 10 | 15 | 20 | 40 | 60 | 70 | 120 | 200
| Standard Pressure (SP) (no vacuum) | 2 | 5 | 10 | 15 | 20 | 40 | 60 | 70 | 120 | 200
| High Pressure (HP) | 2 | 5 | 10 | 15 | 20 | 40 | 60 | 70 | 120 | 200
2E2-S | Low Pressure (LP) (with vacuum) | 4 | 10 | 20 | 28 | 40 | 80 | 120 | 140 | 240 | 400
| Low Pressure (LP) (no vacuum) | 4 | 10 | 20 | 28 | 40 | 80 | 120 | 140 | 240 | 400
| Standard Pressure (SP) (with vacuum) | 4 | 10 | 20 | 28 | 40 | 80 | 120 | 140 | 240 | 400
| Standard Pressure (SP) (no vacuum) | 4 | 10 | 20 | 28 | 40 | 80 | 120 | 140 | 240 | 400
| High Pressure (HP) | 4 | 10 | 20 | 28 | 40 | 80 | 120 | 140 | 240 | 400
5E2-S | Low Pressure (LP) (with vacuum) | 10 | 25 | 50 | 67 | 100 | 200 | 300 | 350 | 600 | 1000
| Low Pressure (LP) (no vacuum) | 10 | 25 | 50 | 67 | 100 | 200 | 300 | 350 | 600 | 1000
| Standard Pressure (SP) (with vacuum) | 10 | 25 | 50 | 67 | 100 | 200 | 300 | 350 | 600 | 1000
| Standard Pressure (SP) (no vacuum) | 10 | 25 | 50 | 67 | 100 | 200 | 300 | 350 | 600 | 1000
| High Pressure (HP) | 10 | 25 | 50 | 67 | 100 | 200 | 300 | 350 | 600 | 1000
1E3-S | Low Pressure (LP) (with vacuum) | 20 | 50 | 100 | 129 | 200 | 400 | 600 | 700 | 1200 | 2000
| Low Pressure (LP) (no vacuum) | 20 | 50 | 100 | 129 | 200 | 400 | 600 | 700 | 1200 | 2000
| Standard Pressure (SP) (with vacuum) | 20 | 50 | 100 | 129 | 200 | 400 | 600 | 700 | 1200 | 2000
| Standard Pressure (SP) (no vacuum) | 20 | 50 | 100 | 129 | 200 | 400 | 600 | 700 | 1200 | 2000
| High Pressure (HP) | 20 | 50 | 100 | 129 | 200 | 400 | 600 | 700 | 1200 | 2000
2E3-S | Low Pressure (LP) (with vacuum) | 40 | 100 | 200 | 248 | 400 | 800 | 1200 | 1400 | 2400 | 4000
| Low Pressure (LP) (no vacuum) | 40 | 100 | 200 | 248 | 400 | 800 | 1200 | 1400 | 2400 | 4000
| Standard Pressure (SP) (with vacuum) | 40 | 100 | 200 | 248 | 400 | 800 | 1200 | 1400 | 2400 | 4000
| Standard Pressure (SP) (no vacuum) | 40 | 100 | 200 | 248 | 400 | 800 | 1200 | 1400 | 2400 | 4000
| High Pressure (HP) | 40 | 100 | 200 | 248 | 400 | 800 | 1200 | 1400 | 2400 | 4000
5E3-S | Low Pressure (LP) (with vacuum) | 100 | 250 | 500 | 596 | 1000 | 2000 | 3000 | 3500 | 6000 | 10000
| Low Pressure (LP) (no vacuum) | 100 | 250 | 500 | 596 | 1000 | 2000 | 3000 | 3500 | 6000 | 10000
| Standard Pressure (SP) (with vacuum) | 100 | 250 | 500 | 596 | 1000 | 2000 | 3000 | 3500 | 6000 | 10000
| Standard Pressure (SP) (no vacuum) | 100 | 250 | 500 | 596 | 1000 | 2000 | 3000 | 3500 | 6000 | 10000
| High Pressure (HP) | 100 | 250 | 500 | 596 | 1000 | 2000 | 3000 | 3500 | 6000 | 10000
1E4-S | Low Pressure (LP) (with vacuum) | 200 | 500 | 1000 | 1173 | 2000 | 4000 | 6000 | 7000 | 12000 | 20000
| Low Pressure (LP) (no vacuum) | 200 | 500 | 1000 | 1173 | 2000 | 4000 | 6000 | 7000 | 12000 | 20000
| Standard Pressure (SP) (with vacuum) | 200 | 500 | 1000 | 1173 | 2000 | 4000 | 6000 | 7000 | 12000 | 20000
| Standard Pressure (SP) (no vacuum) | 200 | 500 | 1000 | 1173 | 2000 | 4000 | 6000 | 7000 | 12000 | 20000
| High Pressure (HP) | 200 | 500 | 1000 | 1173 | 2000 | 4000 | 6000 | 7000 | 12000 | 20000
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.
Air Ratio = 0.983 | Air | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.5 | 1 | 2 | 2 | 3.95 | 5.9 | 6.9 | 11.8 | 19.7
| Low Pressure (LP) (no vacuum) | 0.2 | 0.5 | 1 | 2 | 2 | 3.95 | 5.9 | 6.9 | 11.8 | 19.7
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.5 | 1 | 2 | 2 | 3.95 | 5.9 | 6.9 | 11.8 | 19.7
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.5 | 1 | 2 | 2 | 3.95 | 5.9 | 6.9 | 11.8 | 19.7
| High Pressure (HP) | 0.2 | 0.5 | 1 | 2 | 2 | 3.95 | 5.9 | 6.9 | 11.8 | 19.7
2E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 1 | 2 | 3.4 | 3.9 | 7.85 | 11.8 | 13.8 | 23.6 | 39.3
| Low Pressure (LP) (no vacuum) | 0.4 | 1 | 2 | 3.4 | 3.9 | 7.85 | 11.8 | 13.8 | 23.6 | 39.3
| Standard Pressure (SP) (with vacuum) | 0.4 | 1 | 2 | 3.4 | 3.9 | 7.85 | 11.8 | 13.8 | 23.6 | 39.3
| Standard Pressure (SP) (no vacuum) | 0.4 | 1 | 2 | 3.4 | 3.9 | 7.85 | 11.8 | 13.8 | 23.6 | 39.3
| High Pressure (HP) | 0.4 | 1 | 2 | 3.4 | 3.9 | 7.85 | 11.8 | 13.8 | 23.6 | 39.3
5E1-S | Low Pressure (LP) (with vacuum) | 1 | 2.5 | 4.9 | 7.6 | 9.8 | 19.65 | 29.5 | 34.4 | 59 | 98.3
| Low Pressure (LP) (no vacuum) | 1 | 2.5 | 4.9 | 7.6 | 9.8 | 19.65 | 29.5 | 34.4 | 59 | 98.3
| Standard Pressure (SP) (with vacuum) | 1 | 2.5 | 4.9 | 7.6 | 9.8 | 19.65 | 29.5 | 34.4 | 59 | 98.3
| Standard Pressure (SP) (no vacuum) | 1 | 2.5 | 4.9 | 7.6 | 9.8 | 19.65 | 29.5 | 34.4 | 59 | 98.3
| High Pressure (HP) | 1 | 2.5 | 4.9 | 7.6 | 9.8 | 19.65 | 29.5 | 34.4 | 59 | 98.3
1E2-S | Low Pressure (LP) (with vacuum) | 2 | 4.9 | 9.8 | 15.2 | 19.7 | 39.35 | 59 | 68.8 | 118 | 196.7
| Low Pressure (LP) (no vacuum) | 2 | 4.9 | 9.8 | 15.2 | 19.7 | 39.35 | 59 | 68.8 | 118 | 196.7
| Standard Pressure (SP) (with vacuum) | 2 | 4.9 | 9.8 | 15.2 | 19.7 | 39.35 | 59 | 68.8 | 118 | 196.7
| Standard Pressure (SP) (no vacuum) | 2 | 4.9 | 9.8 | 15.2 | 19.7 | 39.35 | 59 | 68.8 | 118 | 196.7
| High Pressure (HP) | 2 | 4.9 | 9.8 | 15.2 | 19.7 | 39.35 | 59 | 68.8 | 118 | 196.7
2E2-S | Low Pressure (LP) (with vacuum) | 3.9 | 9.8 | 19.7 | 27.4 | 39.3 | 78.65 | 118 | 137.7 | 236 | 393.4
| Low Pressure (LP) (no vacuum) | 3.9 | 9.8 | 19.7 | 27.4 | 39.3 | 78.65 | 118 | 137.7 | 236 | 393.4
| Standard Pressure (SP) (with vacuum) | 3.9 | 9.8 | 19.7 | 27.4 | 39.3 | 78.65 | 118 | 137.7 | 236 | 393.4
| Standard Pressure (SP) (no vacuum) | 3.9 | 9.8 | 19.7 | 27.4 | 39.3 | 78.65 | 118 | 137.7 | 236 | 393.4
| High Pressure (HP) | 3.9 | 9.8 | 19.7 | 27.4 | 39.3 | 78.65 | 118 | 137.7 | 236 | 393.4
5E2-S | Low Pressure (LP) (with vacuum) | 9.8 | 24.6 | 49.2 | 67.1 | 98.3 | 196.675 | 295.05 | 344.2 | 590.1 | 983.5
| Low Pressure (LP) (no vacuum) | 9.8 | 24.6 | 49.2 | 67.1 | 98.3 | 196.675 | 295.05 | 344.2 | 590.1 | 983.5
| Standard Pressure (SP) (with vacuum) | 9.8 | 24.6 | 49.2 | 67.1 | 98.3 | 196.675 | 295.05 | 344.2 | 590.1 | 983.5
| Standard Pressure (SP) (no vacuum) | 9.8 | 24.6 | 49.2 | 67.1 | 98.3 | 196.675 | 295.05 | 344.2 | 590.1 | 983.5
| High Pressure (HP) | 9.8 | 24.6 | 49.2 | 67.1 | 98.3 | 196.675 | 295.05 | 344.2 | 590.1 | 983.5
1E3-S | Low Pressure (LP) (with vacuum) | 19.7 | 49.2 | 98.3 | 126.7 | 196.7 | 393.375 | 590.05 | 688.4 | 1180.2 | 1967
| Low Pressure (LP) (no vacuum) | 19.7 | 49.2 | 98.3 | 126.7 | 196.7 | 393.375 | 590.05 | 688.4 | 1180.2 | 1967
| Standard Pressure (SP) (with vacuum) | 19.7 | 49.2 | 98.3 | 126.7 | 196.7 | 393.375 | 590.05 | 688.4 | 1180.2 | 1967
| Standard Pressure (SP) (no vacuum) | 19.7 | 49.2 | 98.3 | 126.7 | 196.7 | 393.375 | 590.05 | 688.4 | 1180.2 | 1967
| High Pressure (HP) | 19.7 | 49.2 | 98.3 | 126.7 | 196.7 | 393.375 | 590.05 | 688.4 | 1180.2 | 1967
2E3-S | Low Pressure (LP) (with vacuum) | 39.3 | 98.3 | 196.7 | 243.9 | 393.4 | 786.8 | 1180.2 | 1376.9 | 2360.4 | 3933.9
| Low Pressure (LP) (no vacuum) | 39.3 | 98.3 | 196.7 | 243.9 | 393.4 | 786.8 | 1180.2 | 1376.9 | 2360.4 | 3933.9
| Standard Pressure (SP) (with vacuum) | 39.3 | 98.3 | 196.7 | 243.9 | 393.4 | 786.8 | 1180.2 | 1376.9 | 2360.4 | 3933.9
| Standard Pressure (SP) (no vacuum) | 39.3 | 98.3 | 196.7 | 243.9 | 393.4 | 786.8 | 1180.2 | 1376.9 | 2360.4 | 3933.9
| High Pressure (HP) | 39.3 | 98.3 | 196.7 | 243.9 | 393.4 | 786.8 | 1180.2 | 1376.9 | 2360.4 | 3933.9
5E3-S | Low Pressure (LP) (with vacuum) | 98.3 | 245.9 | 491.7 | 585.2 | 983.5 | 1966.975 | 2950.45 | 3442.2 | 5900.9 | 9834.8
| Low Pressure (LP) (no vacuum) | 98.3 | 245.9 | 491.7 | 585.2 | 983.5 | 1966.975 | 2950.45 | 3442.2 | 5900.9 | 9834.8
| Standard Pressure (SP) (with vacuum) | 98.3 | 245.9 | 491.7 | 585.2 | 983.5 | 1966.975 | 2950.45 | 3442.2 | 5900.9 | 9834.8
| Standard Pressure (SP) (no vacuum) | 98.3 | 245.9 | 491.7 | 585.2 | 983.5 | 1966.975 | 2950.45 | 3442.2 | 5900.9 | 9834.8
| High Pressure (HP) | 98.3 | 245.9 | 491.7 | 585.2 | 983.5 | 1966.975 | 2950.45 | 3442.2 | 5900.9 | 9834.8
1E4-S | Low Pressure (LP) (with vacuum) | 196.7 | 491.7 | 983.5 | 1170.3 | 1967 | 3933.95 | 5900.9 | 6884.4 | 11801.8 | 19669.7
| Low Pressure (LP) (no vacuum) | 196.7 | 491.7 | 983.5 | 1170.3 | 1967 | 3933.95 | 5900.9 | 6884.4 | 11801.8 | 19669.7
| Standard Pressure (SP) (with vacuum) | 196.7 | 491.7 | 983.5 | 1170.3 | 1967 | 3933.95 | 5900.9 | 6884.4 | 11801.8 | 19669.7
| Standard Pressure (SP) (no vacuum) | 196.7 | 491.7 | 983.5 | 1170.3 | 1967 | 3933.95 | 5900.9 | 6884.4 | 11801.8 | 19669.7
| High Pressure (HP) | 196.7 | 491.7 | 983.5 | 1170.3 | 1967 | 3933.95 | 5900.9 | 6884.4 | 11801.8 | 19669.7
[1] Flow values in table are valid only when critical flow is established [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path marching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.
Argon Ratio = 0.837 | Ar | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.4 | 0.8 | 1.7 | 1.7 | 3.275 | 4.85 | 5.9 | 10 | 16.7
| Low Pressure (LP) (no vacuum) | 0.2 | 0.4 | 0.8 | 1.7 | 1.7 | 3.275 | 4.85 | 5.9 | 10 | 16.7
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.4 | 0.8 | 1.7 | 1.7 | 3.275 | 4.85 | 5.9 | 10 | 16.7
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.4 | 0.8 | 1.7 | 1.7 | 3.275 | 4.85 | 5.9 | 10 | 16.7
| High Pressure (HP) | 0.2 | 0.4 | 0.8 | 1.7 | 1.7 | 3.275 | 4.85 | 5.9 | 10 | 16.7
2E1-S | Low Pressure (LP) (with vacuum) | 0.3 | 0.8 | 1.7 | 3 | 3.3 | 6.2438 | 9.1875 | 11.7 | 20.1 | 33.5
| Low Pressure (LP) (no vacuum) | 0.3 | 0.8 | 1.7 | 3 | 3.3 | 6.2438 | 9.1875 | 11.7 | 20.1 | 33.5
| Standard Pressure (SP) (with vacuum) | 0.3 | 0.8 | 1.7 | 3 | 3.3 | 6.2438 | 9.1875 | 11.7 | 20.1 | 33.5
| Standard Pressure (SP) (no vacuum) | 0.3 | 0.8 | 1.7 | 3 | 3.3 | 6.2438 | 9.1875 | 11.7 | 20.1 | 33.5
| High Pressure (HP) | 0.3 | 0.8 | 1.7 | 3 | 3.3 | 6.2438 | 9.1875 | 11.7 | 20.1 | 33.5
5E1-S | Low Pressure (LP) (with vacuum) | 0.8 | 2.1 | 4.2 | 6.5 | 8.4 | 15.7125 | 23.025 | 29.3 | 50.2 | 83.7
| Low Pressure (LP) (no vacuum) | 0.8 | 2.1 | 4.2 | 6.5 | 8.4 | 15.7125 | 23.025 | 29.3 | 50.2 | 83.7
| Standard Pressure (SP) (with vacuum) | 0.8 | 2.1 | 4.2 | 6.5 | 8.4 | 15.7125 | 23.025 | 29.3 | 50.2 | 83.7
| Standard Pressure (SP) (no vacuum) | 0.8 | 2.1 | 4.2 | 6.5 | 8.4 | 15.7125 | 23.025 | 29.3 | 50.2 | 83.7
| High Pressure (HP) | 0.8 | 2.1 | 4.2 | 6.5 | 8.4 | 15.7125 | 23.025 | 29.3 | 50.2 | 83.7
1E2-S | Low Pressure (LP) (with vacuum) | 1.7 | 4.2 | 8.4 | 12.9 | 16.7 | 31.3687 | 46.0375 | 58.6 | 100.5 | 167.4
| Low Pressure (LP) (no vacuum) | 1.7 | 4.2 | 8.4 | 12.9 | 16.7 | 31.3687 | 46.0375 | 58.6 | 100.5 | 167.4
| Standard Pressure (SP) (with vacuum) | 1.7 | 4.2 | 8.4 | 12.9 | 16.7 | 31.3687 | 46.0375 | 58.6 | 100.5 | 167.4
| Standard Pressure (SP) (no vacuum) | 1.7 | 4.2 | 8.4 | 12.9 | 16.7 | 31.3687 | 46.0375 | 58.6 | 100.5 | 167.4
| High Pressure (HP) | 1.7 | 4.2 | 8.4 | 12.9 | 16.7 | 31.3687 | 46.0375 | 58.6 | 100.5 | 167.4
2E2-S | Low Pressure (LP) (with vacuum) | 3.3 | 8.4 | 16.7 | 23.3 | 33.5 | 62.7938 | 92.0875 | 117.2 | 200.9 | 334.9
| Low Pressure (LP) (no vacuum) | 3.3 | 8.4 | 16.7 | 23.3 | 33.5 | 62.7938 | 92.0875 | 117.2 | 200.9 | 334.9
| Standard Pressure (SP) (with vacuum) | 3.3 | 8.4 | 16.7 | 23.3 | 33.5 | 62.7938 | 92.0875 | 117.2 | 200.9 | 334.9
| Standard Pressure (SP) (no vacuum) | 3.3 | 8.4 | 16.7 | 23.3 | 33.5 | 62.7938 | 92.0875 | 117.2 | 200.9 | 334.9
| High Pressure (HP) | 3.3 | 8.4 | 16.7 | 23.3 | 33.5 | 62.7938 | 92.0875 | 117.2 | 200.9 | 334.9
5E2-S | Low Pressure (LP) (with vacuum) | 8.4 | 20.9 | 41.9 | 57.1 | 83.7 | 156.9563 | 230.2125 | 293 | 502.3 | 837.2
| Low Pressure (LP) (no vacuum) | 8.4 | 20.9 | 41.9 | 57.1 | 83.7 | 156.9563 | 230.2125 | 293 | 502.3 | 837.2
| Standard Pressure (SP) (with vacuum) | 8.4 | 20.9 | 41.9 | 57.1 | 83.7 | 156.9563 | 230.2125 | 293 | 502.3 | 837.2
| Standard Pressure (SP) (no vacuum) | 8.4 | 20.9 | 41.9 | 57.1 | 83.7 | 156.9563 | 230.2125 | 293 | 502.3 | 837.2
| High Pressure (HP) | 8.4 | 20.9 | 41.9 | 57.1 | 83.7 | 156.9563 | 230.2125 | 293 | 502.3 | 837.2
1E3-S | Low Pressure (LP) (with vacuum) | 16.7 | 41.9 | 83.7 | 107.8 | 167.4 | 313.9125 | 460.425 | 586 | 1004.6 | 1674.4
| Low Pressure (LP) (no vacuum) | 16.7 | 41.9 | 83.7 | 107.8 | 167.4 | 313.9125 | 460.425 | 586 | 1004.6 | 1674.4
| Standard Pressure (SP) (with vacuum) | 16.7 | 41.9 | 83.7 | 107.8 | 167.4 | 313.9125 | 460.425 | 586 | 1004.6 | 1674.4
| Standard Pressure (SP) (no vacuum) | 16.7 | 41.9 | 83.7 | 107.8 | 167.4 | 313.9125 | 460.425 | 586 | 1004.6 | 1674.4
| High Pressure (HP) | 16.7 | 41.9 | 83.7 | 107.8 | 167.4 | 313.9125 | 460.425 | 586 | 1004.6 | 1674.4
2E3-S | Low Pressure (LP) (with vacuum) | 33.5 | 83.7 | 167.4 | 207.6 | 334.9 | 627.9188 | 920.9375 | 1172.1 | 2009.3 | 3348.8
| Low Pressure (LP) (no vacuum) | 33.5 | 83.7 | 167.4 | 207.6 | 334.9 | 627.9188 | 920.9375 | 1172.1 | 2009.3 | 3348.8
| Standard Pressure (SP) (with vacuum) | 33.5 | 83.7 | 167.4 | 207.6 | 334.9 | 627.9188 | 920.9375 | 1172.1 | 2009.3 | 3348.8
| Standard Pressure (SP) (no vacuum) | 33.5 | 83.7 | 167.4 | 207.6 | 334.9 | 627.9188 | 920.9375 | 1172.1 | 2009.3 | 3348.8
| High Pressure (HP) | 33.5 | 83.7 | 167.4 | 207.6 | 334.9 | 627.9188 | 920.9375 | 1172.1 | 2009.3 | 3348.8
5E3-S | Low Pressure (LP) (with vacuum) | 83.7 | 209.3 | 418.6 | 498.2 | 837.2 | 1569.75 | 2302.3 | 2930.2 | 5023.2 | 8372
| Low Pressure (LP) (no vacuum) | 83.7 | 209.3 | 418.6 | 498.2 | 837.2 | 1569.75 | 2302.3 | 2930.2 | 5023.2 | 8372
| Standard Pressure (SP) (with vacuum) | 83.7 | 209.3 | 418.6 | 498.2 | 837.2 | 1569.75 | 2302.3 | 2930.2 | 5023.2 | 8372
| Standard Pressure (SP) (no vacuum) | 83.7 | 209.3 | 418.6 | 498.2 | 837.2 | 1569.75 | 2302.3 | 2930.2 | 5023.2 | 8372
| High Pressure (HP) | 83.7 | 209.3 | 418.6 | 498.2 | 837.2 | 1569.75 | 2302.3 | 2930.2 | 5023.2 | 8372
1E4-S | Low Pressure (LP) (with vacuum) | 167.4 | 418.6 | 837.2 | 996.2 | 1674.4 | 3139.5 | 4604.6 | 5860.4 | 10046.4 | 16744
| Low Pressure (LP) (no vacuum) | 167.4 | 418.6 | 837.2 | 996.2 | 1674.4 | 3139.5 | 4604.6 | 5860.4 | 10046.4 | 16744
| Standard Pressure (SP) (with vacuum) | 167.4 | 418.6 | 837.2 | 996.2 | 1674.4 | 3139.5 | 4604.6 | 5860.4 | 10046.4 | 16744
| Standard Pressure (SP) (no vacuum) | 167.4 | 418.6 | 837.2 | 996.2 | 1674.4 | 3139.5 | 4604.6 | 5860.4 | 10046.4 | 16744
| High Pressure (HP) | 167.4 | 418.6 | 837.2 | 996.2 | 1674.4 | 3139.5 | 4604.6 | 5860.4 | 10046.4 | 16744
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path matching tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.
Helium Ratio = 2.647 | He | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.5 | 1.3 | 2.6 | 9.4 | 5.3 | 9.7137 | 14.1273 | 18.5 | 31.8 | 52.9
| Low Pressure (LP) (no vacuum) | 0.5 | 1.3 | 2.6 | 9.4 | 5.3 | 9.7137 | 14.1273 | 18.5 | 31.8 | 52.9
| Standard Pressure (SP) (with vacuum) | 0.5 | 1.3 | 2.6 | 9.4 | 5.3 | 9.7137 | 14.1273 | 18.5 | 31.8 | 52.9
| Standard Pressure (SP) (no vacuum) | 0.5 | 1.3 | 2.6 | 9.4 | 5.3 | 9.7137 | 14.1273 | 18.5 | 31.8 | 52.9
| High Pressure (HP) | 0.5 | 1.3 | 2.6 | 9.4 | 5.3 | 9.7137 | 14.1273 | 18.5 | 31.8 | 52.9
2E1-S | Low Pressure (LP) (with vacuum) | 1.1 | 2.6 | 5.3 | 13.1 | 10.6 | 21.2 | 29.15 | 37.1 | 63.5 | 105.9
| Low Pressure (LP) (no vacuum) | 1.1 | 2.6 | 5.3 | 13.1 | 10.6 | 21.2 | 29.15 | 37.1 | 63.5 | 105.9
| Standard Pressure (SP) (with vacuum) | 1.1 | 2.6 | 5.3 | 13.1 | 10.6 | 21.2 | 29.15 | 37.1 | 63.5 | 105.9
| Standard Pressure (SP) (no vacuum) | 1.1 | 2.6 | 5.3 | 13.1 | 10.6 | 21.2 | 29.15 | 37.1 | 63.5 | 105.9
| High Pressure (HP) | 1.1 | 2.6 | 5.3 | 13.1 | 10.6 | 21.2 | 29.15 | 37.1 | 63.5 | 105.9
5E1-S | Low Pressure (LP) (with vacuum) | 2.6 | 6.6 | 13.2 | 25.7 | 26.5 | 52.95 | 72.775 | 92.6 | 158.8 | 264.7
| Low Pressure (LP) (no vacuum) | 2.6 | 6.6 | 13.2 | 25.7 | 26.5 | 52.95 | 72.775 | 92.6 | 158.8 | 264.7
| Standard Pressure (SP) (with vacuum) | 2.6 | 6.6 | 13.2 | 25.7 | 26.5 | 52.95 | 72.775 | 92.6 | 158.8 | 264.7
| Standard Pressure (SP) (no vacuum) | 2.6 | 6.6 | 13.2 | 25.7 | 26.5 | 52.95 | 72.775 | 92.6 | 158.8 | 264.7
| High Pressure (HP) | 2.6 | 6.6 | 13.2 | 25.7 | 26.5 | 52.95 | 72.775 | 92.6 | 158.8 | 264.7
1E2-S | Low Pressure (LP) (with vacuum) | 5.3 | 13.2 | 26.5 | 51.4 | 52.9 | 105.85 | 145.575 | 185.3 | 317.6 | 529.4
| Low Pressure (LP) (no vacuum) | 5.3 | 13.2 | 26.5 | 51.4 | 52.9 | 105.85 | 145.575 | 185.3 | 317.6 | 529.4
| Standard Pressure (SP) (with vacuum) | 5.3 | 13.2 | 26.5 | 51.4 | 52.9 | 105.85 | 145.575 | 185.3 | 317.6 | 529.4
| Standard Pressure (SP) (no vacuum) | 5.3 | 13.2 | 26.5 | 51.4 | 52.9 | 105.85 | 145.575 | 185.3 | 317.6 | 529.4
| High Pressure (HP) | 5.3 | 13.2 | 26.5 | 51.4 | 52.9 | 105.85 | 145.575 | 185.3 | 317.6 | 529.4
2E2-S | Low Pressure (LP) (with vacuum) | 10.6 | 26.5 | 52.9 | 91.5 | 105.9 | 211.775 | 291.1875 | 370.6 | 635.2 | 1058.7
| Low Pressure (LP) (no vacuum) | 10.6 | 26.5 | 52.9 | 91.5 | 105.9 | 211.775 | 291.1875 | 370.6 | 635.2 | 1058.7
| Standard Pressure (SP) (with vacuum) | 10.6 | 26.5 | 52.9 | 91.5 | 105.9 | 211.775 | 291.1875 | 370.6 | 635.2 | 1058.7
| Standard Pressure (SP) (no vacuum) | 10.6 | 26.5 | 52.9 | 91.5 | 105.9 | 211.775 | 291.1875 | 370.6 | 635.2 | 1058.7
| High Pressure (HP) | 10.6 | 26.5 | 52.9 | 91.5 | 105.9 | 211.775 | 291.1875 | 370.6 | 635.2 | 1058.7
5E2-S | Low Pressure (LP) (with vacuum) | 26.5 | 66.2 | 132.3 | 199.4 | 264.7 | 529.375 | 727.8875 | 926.4 | 1588.1 | 2646.8
| Low Pressure (LP) (no vacuum) | 26.5 | 66.2 | 132.3 | 199.4 | 264.7 | 529.375 | 727.8875 | 926.4 | 1588.1 | 2646.8
| Standard Pressure (SP) (with vacuum) | 26.5 | 66.2 | 132.3 | 199.4 | 264.7 | 529.375 | 727.8875 | 926.4 | 1588.1 | 2646.8
| Standard Pressure (SP) (no vacuum) | 26.5 | 66.2 | 132.3 | 199.4 | 264.7 | 529.375 | 727.8875 | 926.4 | 1588.1 | 2646.8
| High Pressure (HP) | 26.5 | 66.2 | 132.3 | 199.4 | 264.7 | 529.375 | 727.8875 | 926.4 | 1588.1 | 2646.8
1E3-S | Low Pressure (LP) (with vacuum) | 52.9 | 132.3 | 264.7 | 398.7 | 529.4 | 1058.75 | 1455.775 | 1852.8 | 3176.2 | 5293.6
| Low Pressure (LP) (no vacuum) | 52.9 | 132.3 | 264.7 | 398.7 | 529.4 | 1058.75 | 1455.775 | 1852.8 | 3176.2 | 5293.6
| Standard Pressure (SP) (with vacuum) | 52.9 | 132.3 | 264.7 | 398.7 | 529.4 | 1058.75 | 1455.775 | 1852.8 | 3176.2 | 5293.6
| Standard Pressure (SP) (no vacuum) | 52.9 | 132.3 | 264.7 | 398.7 | 529.4 | 1058.75 | 1455.775 | 1852.8 | 3176.2 | 5293.6
| High Pressure (HP) | 52.9 | 132.3 | 264.7 | 398.7 | 529.4 | 1058.75 | 1455.775 | 1852.8 | 3176.2 | 5293.6
2E3-S | Low Pressure (LP) (with vacuum) | 105.9 | 264.7 | 529.4 | 695.1 | 1058.7 | 2117.425 | 2911.4625 | 3705.5 | 6352.3 | 10587.2
| Low Pressure (LP) (no vacuum) | 105.9 | 264.7 | 529.4 | 695.1 | 1058.7 | 2117.425 | 2911.4625 | 3705.5 | 6352.3 | 10587.2
| Standard Pressure (SP) (with vacuum) | 105.9 | 264.7 | 529.4 | 695.1 | 1058.7 | 2117.425 | 2911.4625 | 3705.5 | 6352.3 | 10587.2
| Standard Pressure (SP) (no vacuum) | 105.9 | 264.7 | 529.4 | 695.1 | 1058.7 | 2117.425 | 2911.4625 | 3705.5 | 6352.3 | 10587.2
| High Pressure (HP) | 105.9 | 264.7 | 529.4 | 695.1 | 1058.7 | 2117.425 | 2911.4625 | 3705.5 | 6352.3 | 10587.2
5E3-S | Low Pressure (LP) (with vacuum) | 264.7 | 661.7 | 1323.4 | 1737.8 | 2646.8 | 5293.6 | 7278.7 | 9263.8 | 15880.9 | 26468.1
| Low Pressure (LP) (no vacuum) | 264.7 | 661.7 | 1323.4 | 1737.8 | 2646.8 | 5293.6 | 7278.7 | 9263.8 | 15880.9 | 26468.1
| Standard Pressure (SP) (with vacuum) | 264.7 | 661.7 | 1323.4 | 1737.8 | 2646.8 | 5293.6 | 7278.7 | 9263.8 | 15880.9 | 26468.1
| Standard Pressure (SP) (no vacuum) | 264.7 | 661.7 | 1323.4 | 1737.8 | 2646.8 | 5293.6 | 7278.7 | 9263.8 | 15880.9 | 26468.1
| High Pressure (HP) | 264.7 | 661.7 | 1323.4 | 1737.8 | 2646.8 | 5293.6 | 7278.7 | 9263.8 | 15880.9 | 26468.1
1E4-S | Low Pressure (LP) (with vacuum) | 529.4 | 1323.4 | 2646.8 | 3281 | 5293.6 | 10587.225 | 14557.4625 | 18527.7 | 31761.7 | 52936.2
| Low Pressure (LP) (no vacuum) | 529.4 | 1323.4 | 2646.8 | 3281 | 5293.6 | 10587.225 | 14557.4625 | 18527.7 | 31761.7 | 52936.2
| Standard Pressure (SP) (with vacuum) | 529.4 | 1323.4 | 2646.8 | 3281 | 5293.6 | 10587.225 | 14557.4625 | 18527.7 | 31761.7 | 52936.2
| Standard Pressure (SP) (no vacuum) | 529.4 | 1323.4 | 2646.8 | 3281 | 5293.6 | 10587.225 | 14557.4625 | 18527.7 | 31761.7 | 52936.2
| High Pressure (HP) | 529.4 | 1323.4 | 2646.8 | 3281 | 5293.6 | 10587.225 | 14557.4625 | 18527.7 | 31761.7 | 52936.2
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.
Carbon Dioxide Ratio = 0.795 | CO2 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 |  | n/a[4]
| High Pressure (HP) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.05 | 238.6 | 278.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.05 | 238.6 | 278.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.05 | 238.6 | 278.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.05 | 238.6 | 278.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.05 | 238.6 | 278.4 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 15.9 | 39.8 | 79.5 | 102.4 | 159.1 | 318.15 | 477.2 | 556.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 15.9 | 39.8 | 79.5 | 102.4 | 159.1 | 318.15 | 477.2 | 556.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 15.9 | 39.8 | 79.5 | 102.4 | 159.1 | 318.15 | 477.2 | 556.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 15.9 | 39.8 | 79.5 | 102.4 | 159.1 | 318.15 | 477.2 | 556.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 15.9 | 39.8 | 79.5 | 102.4 | 159.1 | 318.15 | 477.2 | 556.7 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 31.8 | 79.5 | 159.1 | 189.3 | 318.1 | 636.225 | 954.35 | 1113.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 31.8 | 79.5 | 159.1 | 189.3 | 318.1 | 636.225 | 954.35 | 1113.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 31.8 | 79.5 | 159.1 | 189.3 | 318.1 | 636.225 | 954.35 | 1113.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 31.8 | 79.5 | 159.1 | 189.3 | 318.1 | 636.225 | 954.35 | 1113.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 31.8 | 79.5 | 159.1 | 189.3 | 318.1 | 636.225 | 954.35 | 1113.4 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 79.5 | 198.8 | 397.7 | 473.2 | 795.3 | 1590.625 | 2385.95 | 2783.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 79.5 | 198.8 | 397.7 | 473.2 | 795.3 | 1590.625 | 2385.95 | 2783.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 79.5 | 198.8 | 397.7 | 473.2 | 795.3 | 1590.625 | 2385.95 | 2783.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 79.5 | 198.8 | 397.7 | 473.2 | 795.3 | 1590.625 | 2385.95 | 2783.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 79.5 | 198.8 | 397.7 | 473.2 | 795.3 | 1590.625 | 2385.95 | 2783.6 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 159.1 | 397.7 | 795.3 | 914.1 | 1590.6 | 3181.25 | 4771.9 | 5567.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 159.1 | 397.7 | 795.3 | 914.1 | 1590.6 | 3181.25 | 4771.9 | 5567.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 159.1 | 397.7 | 795.3 | 914.1 | 1590.6 | 3181.25 | 4771.9 | 5567.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 159.1 | 397.7 | 795.3 | 914.1 | 1590.6 | 3181.25 | 4771.9 | 5567.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 159.1 | 397.7 | 795.3 | 914.1 | 1590.6 | 3181.25 | 4771.9 | 5567.2 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Sulfur Hexafluoride Ratio = 0.435 | SF6 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.75 | 2.6 | 3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.75 | 2.6 | 3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.75 | 2.6 | 3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.75 | 2.6 | 3 |  | n/a[4]
| High Pressure (HP) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.75 | 2.6 | 3 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.4 | 0.9 | 1.4 | 1.7 | 3.45 | 5.2 | 6.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.2 | 0.4 | 0.9 | 1.4 | 1.7 | 3.45 | 5.2 | 6.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.4 | 0.9 | 1.4 | 1.7 | 3.45 | 5.2 | 6.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.4 | 0.9 | 1.4 | 1.7 | 3.45 | 5.2 | 6.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.2 | 0.4 | 0.9 | 1.4 | 1.7 | 3.45 | 5.2 | 6.1 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 1.1 | 2.2 | 3.1 | 4.3 | 8.675 | 13.05 | 15.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.4 | 1.1 | 2.2 | 3.1 | 4.3 | 8.675 | 13.05 | 15.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.4 | 1.1 | 2.2 | 3.1 | 4.3 | 8.675 | 13.05 | 15.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.4 | 1.1 | 2.2 | 3.1 | 4.3 | 8.675 | 13.05 | 15.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.4 | 1.1 | 2.2 | 3.1 | 4.3 | 8.675 | 13.05 | 15.2 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 0.9 | 2.2 | 4.3 | 5.9 | 8.7 | 17.375 | 26.05 | 30.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.9 | 2.2 | 4.3 | 5.9 | 8.7 | 17.375 | 26.05 | 30.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.9 | 2.2 | 4.3 | 5.9 | 8.7 | 17.375 | 26.05 | 30.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.9 | 2.2 | 4.3 | 5.9 | 8.7 | 17.375 | 26.05 | 30.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.9 | 2.2 | 4.3 | 5.9 | 8.7 | 17.375 | 26.05 | 30.4 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 1.7 | 4.3 | 8.7 | 11.4 | 17.4 | 34.8 | 52.2 | 60.9 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.7 | 4.3 | 8.7 | 11.4 | 17.4 | 34.8 | 52.2 | 60.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.7 | 4.3 | 8.7 | 11.4 | 17.4 | 34.8 | 52.2 | 60.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.7 | 4.3 | 8.7 | 11.4 | 17.4 | 34.8 | 52.2 | 60.9 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.7 | 4.3 | 8.7 | 11.4 | 17.4 | 34.8 | 52.2 | 60.9 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 4.3 | 10.9 | 21.7 | 26.9 | 43.5 | 86.975 | 130.45 | 152.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 4.3 | 10.9 | 21.7 | 26.9 | 43.5 | 86.975 | 130.45 | 152.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 4.3 | 10.9 | 21.7 | 26.9 | 43.5 | 86.975 | 130.45 | 152.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 4.3 | 10.9 | 21.7 | 26.9 | 43.5 | 86.975 | 130.45 | 152.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 4.3 | 10.9 | 21.7 | 26.9 | 43.5 | 86.975 | 130.45 | 152.2 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 8.7 | 21.7 | 43.5 | 53.8 | 87 | 173.925 | 260.85 | 304.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 8.7 | 21.7 | 43.5 | 53.8 | 87 | 173.925 | 260.85 | 304.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 8.7 | 21.7 | 43.5 | 53.8 | 87 | 173.925 | 260.85 | 304.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 8.7 | 21.7 | 43.5 | 53.8 | 87 | 173.925 | 260.85 | 304.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 8.7 | 21.7 | 43.5 | 53.8 | 87 | 173.925 | 260.85 | 304.3 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 17.4 | 43.5 | 87 | 100 | 173.9 | 347.825 | 521.75 | 608.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 17.4 | 43.5 | 87 | 100 | 173.9 | 347.825 | 521.75 | 608.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 17.4 | 43.5 | 87 | 100 | 173.9 | 347.825 | 521.75 | 608.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 17.4 | 43.5 | 87 | 100 | 173.9 | 347.825 | 521.75 | 608.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 17.4 | 43.5 | 87 | 100 | 173.9 | 347.825 | 521.75 | 608.7 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 43.5 | 108.7 | 217.4 | 249.9 | 434.8 | 869.55 | 1304.3 | 1521.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 43.5 | 108.7 | 217.4 | 249.9 | 434.8 | 869.55 | 1304.3 | 1521.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 43.5 | 108.7 | 217.4 | 249.9 | 434.8 | 869.55 | 1304.3 | 1521.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 43.5 | 108.7 | 217.4 | 249.9 | 434.8 | 869.55 | 1304.3 | 1521.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 43.5 | 108.7 | 217.4 | 249.9 | 434.8 | 869.55 | 1304.3 | 1521.7 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 87 | 217.4 | 434.8 | 499.7 | 869.5 | 1739.05 | 2608.6 | 3043.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 87 | 217.4 | 434.8 | 499.7 | 869.5 | 1739.05 | 2608.6 | 3043.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 87 | 217.4 | 434.8 | 499.7 | 869.5 | 1739.05 | 2608.6 | 3043.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 87 | 217.4 | 434.8 | 499.7 | 869.5 | 1739.05 | 2608.6 | 3043.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 87 | 217.4 | 434.8 | 499.7 | 869.5 | 1739.05 | 2608.6 | 3043.4 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Xenon Ratio = 0.460 | Xe | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.7 | 1.6 | 2.5 | 3.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.7 | 1.6 | 2.5 | 3.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.7 | 1.6 | 2.5 | 3.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.7 | 1.6 | 2.5 | 3.2 |  | n/a[4]
| High Pressure (HP) | 0.1 | 0.2 | 0.4 | 0.8 | 0.7 | 1.6 | 2.5 | 3.2 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.4 | 0.7 | 1.4 | 1.5 | 3.25 | 5 | 6.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.4 | 0.7 | 1.4 | 1.5 | 3.25 | 5 | 6.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.4 | 0.7 | 1.4 | 1.5 | 3.25 | 5 | 6.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.4 | 0.7 | 1.4 | 1.5 | 3.25 | 5 | 6.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.1 | 0.4 | 0.7 | 1.4 | 1.5 | 3.25 | 5 | 6.4 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 0.9 | 1.8 | 3.6 | 3.6 | 8.1 | 12.6 | 16.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.4 | 0.9 | 1.8 | 3.6 | 3.6 | 8.1 | 12.6 | 16.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.4 | 0.9 | 1.8 | 3.6 | 3.6 | 8.1 | 12.6 | 16.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.4 | 0.9 | 1.8 | 3.6 | 3.6 | 8.1 | 12.6 | 16.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.4 | 0.9 | 1.8 | 3.6 | 3.6 | 8.1 | 12.6 | 16.1 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 0.7 | 1.8 | 3.6 | 6.5 | 7.3 | 16.25 | 25.2 | 32.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.7 | 1.8 | 3.6 | 6.5 | 7.3 | 16.25 | 25.2 | 32.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.7 | 1.8 | 3.6 | 6.5 | 7.3 | 16.25 | 25.2 | 32.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.7 | 1.8 | 3.6 | 6.5 | 7.3 | 16.25 | 25.2 | 32.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.7 | 1.8 | 3.6 | 6.5 | 7.3 | 16.25 | 25.2 | 32.2 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 1.5 | 3.6 | 7.3 | 12.9 | 14.6 | 32.525 | 50.45 | 64.5 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.5 | 3.6 | 7.3 | 12.9 | 14.6 | 32.525 | 50.45 | 64.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.5 | 3.6 | 7.3 | 12.9 | 14.6 | 32.525 | 50.45 | 64.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.5 | 3.6 | 7.3 | 12.9 | 14.6 | 32.525 | 50.45 | 64.5 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.5 | 3.6 | 7.3 | 12.9 | 14.6 | 32.525 | 50.45 | 64.5 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 3.6 | 9.1 | 18.2 | 29.7 | 36.4 | 81.25 | 126.1 | 161.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 3.6 | 9.1 | 18.2 | 29.7 | 36.4 | 81.25 | 126.1 | 161.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 3.6 | 9.1 | 18.2 | 29.7 | 36.4 | 81.25 | 126.1 | 161.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 3.6 | 9.1 | 18.2 | 29.7 | 36.4 | 81.25 | 126.1 | 161.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 3.6 | 9.1 | 18.2 | 29.7 | 36.4 | 81.25 | 126.1 | 161.2 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 7.3 | 18.2 | 36.4 | 59.3 | 72.8 | 162.475 | 252.15 | 322.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 7.3 | 18.2 | 36.4 | 59.3 | 72.8 | 162.475 | 252.15 | 322.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 7.3 | 18.2 | 36.4 | 59.3 | 72.8 | 162.475 | 252.15 | 322.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 7.3 | 18.2 | 36.4 | 59.3 | 72.8 | 162.475 | 252.15 | 322.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 7.3 | 18.2 | 36.4 | 59.3 | 72.8 | 162.475 | 252.15 | 322.3 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 14.6 | 36.4 | 72.8 | 109.6 | 145.6 | 324.95 | 504.3 | 644.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 14.6 | 36.4 | 72.8 | 109.6 | 145.6 | 324.95 | 504.3 | 644.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 14.6 | 36.4 | 72.8 | 109.6 | 145.6 | 324.95 | 504.3 | 644.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 14.6 | 36.4 | 72.8 | 109.6 | 145.6 | 324.95 | 504.3 | 644.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 14.6 | 36.4 | 72.8 | 109.6 | 145.6 | 324.95 | 504.3 | 644.6 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 36.4 | 91 | 182 | 267.2 | 364 | 812.375 | 1260.75 | 1611.5 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 36.4 | 91 | 182 | 267.2 | 364 | 812.375 | 1260.75 | 1611.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 36.4 | 91 | 182 | 267.2 | 364 | 812.375 | 1260.75 | 1611.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 36.4 | 91 | 182 | 267.2 | 364 | 812.375 | 1260.75 | 1611.5 | n/a[4] | n/a[4]
| High Pressure (HP) | 36.4 | 91 | 182 | 267.2 | 364 | 812.375 | 1260.75 | 1611.5 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 72.8 | 182 | 364 | 529.2 | 728 | 1624.75 | 2521.5 | 3223.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 72.8 | 182 | 364 | 529.2 | 728 | 1624.75 | 2521.5 | 3223.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 72.8 | 182 | 364 | 529.2 | 728 | 1624.75 | 2521.5 | 3223.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 72.8 | 182 | 364 | 529.2 | 728 | 1624.75 | 2521.5 | 3223.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 72.8 | 182 | 364 | 529.2 | 728 | 1624.75 | 2521.5 | 3223.1 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Butane Ratio = 0.680 | C4H10 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.3 | 0.7 | 1 | 1.4 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.3 | 0.7 | 1 | 1.4 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.3 | 0.7 | 1 | 1.4 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.3 | 0.7 | 1 | 1.4 | n/a[8] | n/a[8] |  |  | n/a[4]
| High Pressure (HP) | 0.1 | 0.3 | 0.7 | 1 | 1.4 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.3 | 0.7 | 1.4 | 1.9 | 2.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.3 | 0.7 | 1.4 | 1.9 | 2.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.3 | 0.7 | 1.4 | 1.9 | 2.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.3 | 0.7 | 1.4 | 1.9 | 2.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 0.3 | 0.7 | 1.4 | 1.9 | 2.7 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.7 | 1.7 | 3.4 | 4.8 | 6.8 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.7 | 1.7 | 3.4 | 4.8 | 6.8 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.7 | 1.7 | 3.4 | 4.8 | 6.8 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.7 | 1.7 | 3.4 | 4.8 | 6.8 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 0.7 | 1.7 | 3.4 | 4.8 | 6.8 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 1.4 | 3.4 | 6.8 | 9 | 13.6 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.4 | 3.4 | 6.8 | 9 | 13.6 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.4 | 3.4 | 6.8 | 9 | 13.6 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.4 | 3.4 | 6.8 | 9 | 13.6 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 1.4 | 3.4 | 6.8 | 9 | 13.6 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 2.7 | 6.8 | 13.6 | 18 | 27.2 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 2.7 | 6.8 | 13.6 | 18 | 27.2 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 2.7 | 6.8 | 13.6 | 18 | 27.2 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 2.7 | 6.8 | 13.6 | 18 | 27.2 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 2.7 | 6.8 | 13.6 | 18 | 27.2 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 6.8 | 17 | 34 | 42.2 | 68 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 6.8 | 17 | 34 | 42.2 | 68 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 6.8 | 17 | 34 | 42.2 | 68 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 6.8 | 17 | 34 | 42.2 | 68 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 6.8 | 17 | 34 | 42.2 | 68 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 13.6 | 34 | 68 | 84.4 | 136 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 13.6 | 34 | 68 | 84.4 | 136 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 13.6 | 34 | 68 | 84.4 | 136 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 13.6 | 34 | 68 | 84.4 | 136 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 13.6 | 34 | 68 | 84.4 | 136 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 27.2 | 68 | 136 | 157 | 272 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 27.2 | 68 | 136 | 157 | 272 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 27.2 | 68 | 136 | 157 | 272 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 27.2 | 68 | 136 | 157 | 272 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 27.2 | 68 | 136 | 157 | 272 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 68 | 170 | 340 | 392.4 | 679.9 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 68 | 170 | 340 | 392.4 | 679.9 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 68 | 170 | 340 | 392.4 | 679.9 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 68 | 170 | 340 | 392.4 | 679.9 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 68 | 170 | 340 | 392.4 | 679.9 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 136 | 340 | 679.9 | 784.9 | 1359.8 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 136 | 340 | 679.9 | 784.9 | 1359.8 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 136 | 340 | 679.9 | 784.9 | 1359.8 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 136 | 340 | 679.9 | 784.9 | 1359.8 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 136 | 340 | 679.9 | 784.9 | 1359.8 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Ethane Ratio = 0.960 | C2H6 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.5 | 1 | 1.5 | 1.9 | 3.825 | 5.75 | 6.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.2 | 0.5 | 1 | 1.5 | 1.9 | 3.825 | 5.75 | 6.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.5 | 1 | 1.5 | 1.9 | 3.825 | 5.75 | 6.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.5 | 1 | 1.5 | 1.9 | 3.825 | 5.75 | 6.7 |  | n/a[4]
| High Pressure (HP) | 0.2 | 0.5 | 1 | 1.5 | 1.9 | 3.825 | 5.75 | 6.7 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 1 | 1.9 | 3 | 3.8 | 7.65 | 11.5 | 13.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.4 | 1 | 1.9 | 3 | 3.8 | 7.65 | 11.5 | 13.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.4 | 1 | 1.9 | 3 | 3.8 | 7.65 | 11.5 | 13.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.4 | 1 | 1.9 | 3 | 3.8 | 7.65 | 11.5 | 13.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.4 | 1 | 1.9 | 3 | 3.8 | 7.65 | 11.5 | 13.4 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 1 | 2.4 | 4.8 | 6.7 | 9.6 | 19.2 | 28.8 | 33.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1 | 2.4 | 4.8 | 6.7 | 9.6 | 19.2 | 28.8 | 33.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1 | 2.4 | 4.8 | 6.7 | 9.6 | 19.2 | 28.8 | 33.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1 | 2.4 | 4.8 | 6.7 | 9.6 | 19.2 | 28.8 | 33.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 1 | 2.4 | 4.8 | 6.7 | 9.6 | 19.2 | 28.8 | 33.6 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 1.9 | 4.8 | 9.6 | 13.4 | 19.2 | 38.4 | 57.6 | 67.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.9 | 4.8 | 9.6 | 13.4 | 19.2 | 38.4 | 57.6 | 67.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.9 | 4.8 | 9.6 | 13.4 | 19.2 | 38.4 | 57.6 | 67.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.9 | 4.8 | 9.6 | 13.4 | 19.2 | 38.4 | 57.6 | 67.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.9 | 4.8 | 9.6 | 13.4 | 19.2 | 38.4 | 57.6 | 67.2 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 3.8 | 9.6 | 19.2 | 25.2 | 38.4 | 76.825 | 115.25 | 134.5 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 3.8 | 9.6 | 19.2 | 25.2 | 38.4 | 76.825 | 115.25 | 134.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 3.8 | 9.6 | 19.2 | 25.2 | 38.4 | 76.825 | 115.25 | 134.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 3.8 | 9.6 | 19.2 | 25.2 | 38.4 | 76.825 | 115.25 | 134.5 | n/a[4] | n/a[4]
| High Pressure (HP) | 3.8 | 9.6 | 19.2 | 25.2 | 38.4 | 76.825 | 115.25 | 134.5 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 9.6 | 24 | 48 | 61.9 | 96 | 192.05 | 288.1 | 336.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 9.6 | 24 | 48 | 61.9 | 96 | 192.05 | 288.1 | 336.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 9.6 | 24 | 48 | 61.9 | 96 | 192.05 | 288.1 | 336.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 9.6 | 24 | 48 | 61.9 | 96 | 192.05 | 288.1 | 336.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 9.6 | 24 | 48 | 61.9 | 96 | 192.05 | 288.1 | 336.1 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 19.2 | 48 | 96 | 118.7 | 192.1 | 384.175 | 576.25 | 672.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 19.2 | 48 | 96 | 118.7 | 192.1 | 384.175 | 576.25 | 672.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 19.2 | 48 | 96 | 118.7 | 192.1 | 384.175 | 576.25 | 672.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 19.2 | 48 | 96 | 118.7 | 192.1 | 384.175 | 576.25 | 672.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 19.2 | 48 | 96 | 118.7 | 192.1 | 384.175 | 576.25 | 672.3 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 38.4 | 96 | 192.1 | 228.6 | 384.2 | 768.35 | 1152.5 | 1344.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 38.4 | 96 | 192.1 | 228.6 | 384.2 | 768.35 | 1152.5 | 1344.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 38.4 | 96 | 192.1 | 228.6 | 384.2 | 768.35 | 1152.5 | 1344.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 38.4 | 96 | 192.1 | 228.6 | 384.2 | 768.35 | 1152.5 | 1344.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 38.4 | 96 | 192.1 | 228.6 | 384.2 | 768.35 | 1152.5 | 1344.6 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 96 | 240.1 | 480.2 | 551.9 | 960.4 | 1920.8 | 2881.2 | 3361.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 96 | 240.1 | 480.2 | 551.9 | 960.4 | 1920.8 | 2881.2 | 3361.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 96 | 240.1 | 480.2 | 551.9 | 960.4 | 1920.8 | 2881.2 | 3361.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 96 | 240.1 | 480.2 | 551.9 | 960.4 | 1920.8 | 2881.2 | 3361.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 96 | 240.1 | 480.2 | 551.9 | 960.4 | 1920.8 | 2881.2 | 3361.4 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 192.1 | 480.2 | 960.4 | 1103.8 | 1920.8 | 3841.6 | 5762.4 | 6722.8 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 192.1 | 480.2 | 960.4 | 1103.8 | 1920.8 | 3841.6 | 5762.4 | 6722.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 192.1 | 480.2 | 960.4 | 1103.8 | 1920.8 | 3841.6 | 5762.4 | 6722.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 192.1 | 480.2 | 960.4 | 1103.8 | 1920.8 | 3841.6 | 5762.4 | 6722.8 | n/a[4] | n/a[4]
| High Pressure (HP) | 192.1 | 480.2 | 960.4 | 1103.8 | 1920.8 | 3841.6 | 5762.4 | 6722.8 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Ethylene Ratio = 0.996 | C2H4 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.5 | 1 | 1.7 | 2 | 4 | 6 | 7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.2 | 0.5 | 1 | 1.7 | 2 | 4 | 6 | 7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.5 | 1 | 1.7 | 2 | 4 | 6 | 7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.5 | 1 | 1.7 | 2 | 4 | 6 | 7 |  | n/a[4]
| High Pressure (HP) | 0.2 | 0.5 | 1 | 1.7 | 2 | 4 | 6 | 7 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 1 | 2 | 3.1 | 4 | 7.975 | 11.95 | 13.9 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.4 | 1 | 2 | 3.1 | 4 | 7.975 | 11.95 | 13.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.4 | 1 | 2 | 3.1 | 4 | 7.975 | 11.95 | 13.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.4 | 1 | 2 | 3.1 | 4 | 7.975 | 11.95 | 13.9 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.4 | 1 | 2 | 3.1 | 4 | 7.975 | 11.95 | 13.9 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 1 | 2.5 | 5 | 7.5 | 10 | 19.95 | 29.9 | 34.9 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1 | 2.5 | 5 | 7.5 | 10 | 19.95 | 29.9 | 34.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1 | 2.5 | 5 | 7.5 | 10 | 19.95 | 29.9 | 34.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1 | 2.5 | 5 | 7.5 | 10 | 19.95 | 29.9 | 34.9 | n/a[4] | n/a[4]
| High Pressure (HP) | 1 | 2.5 | 5 | 7.5 | 10 | 19.95 | 29.9 | 34.9 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 2 | 5 | 10 | 13.9 | 19.9 | 39.825 | 59.75 | 69.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 2 | 5 | 10 | 13.9 | 19.9 | 39.825 | 59.75 | 69.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 2 | 5 | 10 | 13.9 | 19.9 | 39.825 | 59.75 | 69.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 2 | 5 | 10 | 13.9 | 19.9 | 39.825 | 59.75 | 69.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 2 | 5 | 10 | 13.9 | 19.9 | 39.825 | 59.75 | 69.7 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 4 | 10 | 19.9 | 27.7 | 39.8 | 79.65 | 119.5 | 139.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 4 | 10 | 19.9 | 27.7 | 39.8 | 79.65 | 119.5 | 139.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 4 | 10 | 19.9 | 27.7 | 39.8 | 79.65 | 119.5 | 139.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 4 | 10 | 19.9 | 27.7 | 39.8 | 79.65 | 119.5 | 139.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 4 | 10 | 19.9 | 27.7 | 39.8 | 79.65 | 119.5 | 139.4 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 10 | 24.9 | 49.8 | 64.2 | 99.6 | 199.15 | 298.7 | 348.5 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 10 | 24.9 | 49.8 | 64.2 | 99.6 | 199.15 | 298.7 | 348.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 10 | 24.9 | 49.8 | 64.2 | 99.6 | 199.15 | 298.7 | 348.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 10 | 24.9 | 49.8 | 64.2 | 99.6 | 199.15 | 298.7 | 348.5 | n/a[4] | n/a[4]
| High Pressure (HP) | 10 | 24.9 | 49.8 | 64.2 | 99.6 | 199.15 | 298.7 | 348.5 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 19.9 | 49.8 | 99.6 | 128.2 | 199.2 | 398.325 | 597.45 | 697 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 19.9 | 49.8 | 99.6 | 128.2 | 199.2 | 398.325 | 597.45 | 697 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 19.9 | 49.8 | 99.6 | 128.2 | 199.2 | 398.325 | 597.45 | 697 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 19.9 | 49.8 | 99.6 | 128.2 | 199.2 | 398.325 | 597.45 | 697 | n/a[4] | n/a[4]
| High Pressure (HP) | 19.9 | 49.8 | 99.6 | 128.2 | 199.2 | 398.325 | 597.45 | 697 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 39.8 | 99.6 | 199.2 | 237 | 398.3 | 796.625 | 1194.95 | 1394.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 39.8 | 99.6 | 199.2 | 237 | 398.3 | 796.625 | 1194.95 | 1394.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 39.8 | 99.6 | 199.2 | 237 | 398.3 | 796.625 | 1194.95 | 1394.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 39.8 | 99.6 | 199.2 | 237 | 398.3 | 796.625 | 1194.95 | 1394.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 39.8 | 99.6 | 199.2 | 237 | 398.3 | 796.625 | 1194.95 | 1394.1 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 99.6 | 248.9 | 497.9 | 572.2 | 995.8 | 1991.575 | 2987.35 | 3485.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 99.6 | 248.9 | 497.9 | 572.2 | 995.8 | 1991.575 | 2987.35 | 3485.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 99.6 | 248.9 | 497.9 | 572.2 | 995.8 | 1991.575 | 2987.35 | 3485.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 99.6 | 248.9 | 497.9 | 572.2 | 995.8 | 1991.575 | 2987.35 | 3485.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 99.6 | 248.9 | 497.9 | 572.2 | 995.8 | 1991.575 | 2987.35 | 3485.2 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 199.2 | 497.9 | 995.8 | 1144.4 | 1991.6 | 3983.15 | 5974.7 | 6970.5 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 199.2 | 497.9 | 995.8 | 1144.4 | 1991.6 | 3983.15 | 5974.7 | 6970.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 199.2 | 497.9 | 995.8 | 1144.4 | 1991.6 | 3983.15 | 5974.7 | 6970.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 199.2 | 497.9 | 995.8 | 1144.4 | 1991.6 | 3983.15 | 5974.7 | 6970.5 | n/a[4] | n/a[4]
| High Pressure (HP) | 199.2 | 497.9 | 995.8 | 1144.4 | 1991.6 | 3983.15 | 5974.7 | 6970.5 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Hydrogen Ratio = 3.730 | H2 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.7 | 1.9 | 3.7 | 8.3 | 7.5 | 14.925 | 22.35 | 26.1 | 44.8 | 74.6
| Low Pressure (LP) (no vacuum) | 0.7 | 1.9 | 3.7 | 8.3 | 7.5 | 14.925 | 22.35 | 26.1 | 44.8 | 74.6
| Standard Pressure (SP) (with vacuum) | 0.7 | 1.9 | 3.7 | 8.3 | 7.5 | 14.925 | 22.35 | 26.1 | 44.8 | 74.6
| Standard Pressure (SP) (no vacuum) | 0.7 | 1.9 | 3.7 | 8.3 | 7.5 | 14.925 | 22.35 | 26.1 | 44.8 | 74.6
| High Pressure (HP) | 0.7 | 1.9 | 3.7 | 8.3 | 7.5 | 14.925 | 22.35 | 26.1 | 44.8 | 74.6
2E1-S | Low Pressure (LP) (with vacuum) | 1.5 | 3.7 | 7.5 | 14.5 | 14.9 | 29.825 | 44.75 | 52.2 | 89.5 | 149.2
| Low Pressure (LP) (no vacuum) | 1.5 | 3.7 | 7.5 | 14.5 | 14.9 | 29.825 | 44.75 | 52.2 | 89.5 | 149.2
| Standard Pressure (SP) (with vacuum) | 1.5 | 3.7 | 7.5 | 14.5 | 14.9 | 29.825 | 44.75 | 52.2 | 89.5 | 149.2
| Standard Pressure (SP) (no vacuum) | 1.5 | 3.7 | 7.5 | 14.5 | 14.9 | 29.825 | 44.75 | 52.2 | 89.5 | 149.2
| High Pressure (HP) | 1.5 | 3.7 | 7.5 | 14.5 | 14.9 | 29.825 | 44.75 | 52.2 | 89.5 | 149.2
5E1-S | Low Pressure (LP) (with vacuum) | 3.7 | 9.3 | 18.6 | 36.2 | 37.3 | 74.575 | 111.85 | 130.5 | 223.8 | 373
| Low Pressure (LP) (no vacuum) | 3.7 | 9.3 | 18.6 | 36.2 | 37.3 | 74.575 | 111.85 | 130.5 | 223.8 | 373
| Standard Pressure (SP) (with vacuum) | 3.7 | 9.3 | 18.6 | 36.2 | 37.3 | 74.575 | 111.85 | 130.5 | 223.8 | 373
| Standard Pressure (SP) (no vacuum) | 3.7 | 9.3 | 18.6 | 36.2 | 37.3 | 74.575 | 111.85 | 130.5 | 223.8 | 373
| High Pressure (HP) | 3.7 | 9.3 | 18.6 | 36.2 | 37.3 | 74.575 | 111.85 | 130.5 | 223.8 | 373
1E2-S | Low Pressure (LP) (with vacuum) | 7.5 | 18.6 | 37.3 | 62.5 | 74.6 | 149.2 | 223.8 | 261.1 | 447.6 | 746
| Low Pressure (LP) (no vacuum) | 7.5 | 18.6 | 37.3 | 62.5 | 74.6 | 149.2 | 223.8 | 261.1 | 447.6 | 746
| Standard Pressure (SP) (with vacuum) | 7.5 | 18.6 | 37.3 | 62.5 | 74.6 | 149.2 | 223.8 | 261.1 | 447.6 | 746
| Standard Pressure (SP) (no vacuum) | 7.5 | 18.6 | 37.3 | 62.5 | 74.6 | 149.2 | 223.8 | 261.1 | 447.6 | 746
| High Pressure (HP) | 7.5 | 18.6 | 37.3 | 62.5 | 74.6 | 149.2 | 223.8 | 261.1 | 447.6 | 746
2E2-S | Low Pressure (LP) (with vacuum) | 14.9 | 37.3 | 74.6 | 114.5 | 149.2 | 298.4 | 447.6 | 522.2 | 895.2 | 1491.9
| Low Pressure (LP) (no vacuum) | 14.9 | 37.3 | 74.6 | 114.5 | 149.2 | 298.4 | 447.6 | 522.2 | 895.2 | 1491.9
| Standard Pressure (SP) (with vacuum) | 14.9 | 37.3 | 74.6 | 114.5 | 149.2 | 298.4 | 447.6 | 522.2 | 895.2 | 1491.9
| Standard Pressure (SP) (no vacuum) | 14.9 | 37.3 | 74.6 | 114.5 | 149.2 | 298.4 | 447.6 | 522.2 | 895.2 | 1491.9
| High Pressure (HP) | 14.9 | 37.3 | 74.6 | 114.5 | 149.2 | 298.4 | 447.6 | 522.2 | 895.2 | 1491.9
5E2-S | Low Pressure (LP) (with vacuum) | 37.3 | 93.2 | 186.5 | 280.9 | 373 | 745.95 | 1118.9 | 1305.4 | 2237.9 | 3729.8
| Low Pressure (LP) (no vacuum) | 37.3 | 93.2 | 186.5 | 280.9 | 373 | 745.95 | 1118.9 | 1305.4 | 2237.9 | 3729.8
| Standard Pressure (SP) (with vacuum) | 37.3 | 93.2 | 186.5 | 280.9 | 373 | 745.95 | 1118.9 | 1305.4 | 2237.9 | 3729.8
| Standard Pressure (SP) (no vacuum) | 37.3 | 93.2 | 186.5 | 280.9 | 373 | 745.95 | 1118.9 | 1305.4 | 2237.9 | 3729.8
| High Pressure (HP) | 37.3 | 93.2 | 186.5 | 280.9 | 373 | 745.95 | 1118.9 | 1305.4 | 2237.9 | 3729.8
1E3-S | Low Pressure (LP) (with vacuum) | 74.6 | 186.5 | 373 | 508.7 | 746 | 1491.95 | 2237.9 | 2610.9 | 4475.8 | 7459.6
| Low Pressure (LP) (no vacuum) | 74.6 | 186.5 | 373 | 508.7 | 746 | 1491.95 | 2237.9 | 2610.9 | 4475.8 | 7459.6
| Standard Pressure (SP) (with vacuum) | 74.6 | 186.5 | 373 | 508.7 | 746 | 1491.95 | 2237.9 | 2610.9 | 4475.8 | 7459.6
| Standard Pressure (SP) (no vacuum) | 74.6 | 186.5 | 373 | 508.7 | 746 | 1491.95 | 2237.9 | 2610.9 | 4475.8 | 7459.6
| High Pressure (HP) | 74.6 | 186.5 | 373 | 508.7 | 746 | 1491.95 | 2237.9 | 2610.9 | 4475.8 | 7459.6
2E3-S | Low Pressure (LP) (with vacuum) | 149.2 | 373 | 746 | 979.6 | 1491.9 | 2983.825 | 4475.75 | 5221.7 | 8951.5 | 14919.2
| Low Pressure (LP) (no vacuum) | 149.2 | 373 | 746 | 979.6 | 1491.9 | 2983.825 | 4475.75 | 5221.7 | 8951.5 | 14919.2
| Standard Pressure (SP) (with vacuum) | 149.2 | 373 | 746 | 979.6 | 1491.9 | 2983.825 | 4475.75 | 5221.7 | 8951.5 | 14919.2
| Standard Pressure (SP) (no vacuum) | 149.2 | 373 | 746 | 979.6 | 1491.9 | 2983.825 | 4475.75 | 5221.7 | 8951.5 | 14919.2
| High Pressure (HP) | 149.2 | 373 | 746 | 979.6 | 1491.9 | 2983.825 | 4475.75 | 5221.7 | 8951.5 | 14919.2
5E3-S | Low Pressure (LP) (with vacuum) | 373 | 932.4 | 1864.9 | 2311.7 | 3729.8 | 7459.6 | 11189.4 | 13054.3 | 22378.8 | 37297.9
| Low Pressure (LP) (no vacuum) | 373 | 932.4 | 1864.9 | 2311.7 | 3729.8 | 7459.6 | 11189.4 | 13054.3 | 22378.8 | 37297.9
| Standard Pressure (SP) (with vacuum) | 373 | 932.4 | 1864.9 | 2311.7 | 3729.8 | 7459.6 | 11189.4 | 13054.3 | 22378.8 | 37297.9
| Standard Pressure (SP) (no vacuum) | 373 | 932.4 | 1864.9 | 2311.7 | 3729.8 | 7459.6 | 11189.4 | 13054.3 | 22378.8 | 37297.9
| High Pressure (HP) | 373 | 932.4 | 1864.9 | 2311.7 | 3729.8 | 7459.6 | 11189.4 | 13054.3 | 22378.8 | 37297.9
1E4-S | Low Pressure (LP) (with vacuum) | 746 | 1864.9 | 3729.8 | 4623.4 | 7459.6 | 14919.2 | 22378.8 | 26108.6 | 44757.5 | 74595.9
| Low Pressure (LP) (no vacuum) | 746 | 1864.9 | 3729.8 | 4623.4 | 7459.6 | 14919.2 | 22378.8 | 26108.6 | 44757.5 | 74595.9
| Standard Pressure (SP) (with vacuum) | 746 | 1864.9 | 3729.8 | 4623.4 | 7459.6 | 14919.2 | 22378.8 | 26108.6 | 44757.5 | 74595.9
| Standard Pressure (SP) (no vacuum) | 746 | 1864.9 | 3729.8 | 4623.4 | 7459.6 | 14919.2 | 22378.8 | 26108.6 | 44757.5 | 74595.9
| High Pressure (HP) | 746 | 1864.9 | 3729.8 | 4623.4 | 7459.6 | 14919.2 | 22378.8 | 26108.6 | 44757.5 | 74595.9
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.
Methane Ratio = 1.320 | CH4 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.3 | 0.7 | 1.3 | 2.6 | 2.6 | 5.25 | 7.9 | 9.2 | 15.8 | 26.4
| Low Pressure (LP) (no vacuum) | 0.3 | 0.7 | 1.3 | 2.6 | 2.6 | 5.25 | 7.9 | 9.2 | 15.8 | 26.4
| Standard Pressure (SP) (with vacuum) | 0.3 | 0.7 | 1.3 | 2.6 | 2.6 | 5.25 | 7.9 | 9.2 | 15.8 | 26.4
| Standard Pressure (SP) (no vacuum) | 0.3 | 0.7 | 1.3 | 2.6 | 2.6 | 5.25 | 7.9 | 9.2 | 15.8 | 26.4
| High Pressure (HP) | 0.3 | 0.7 | 1.3 | 2.6 | 2.6 | 5.25 | 7.9 | 9.2 | 15.8 | 26.4
2E1-S | Low Pressure (LP) (with vacuum) | 0.5 | 1.3 | 2.6 | 4.4 | 5.3 | 10.575 | 15.85 | 18.5 | 31.7 | 52.8
| Low Pressure (LP) (no vacuum) | 0.5 | 1.3 | 2.6 | 4.4 | 5.3 | 10.575 | 15.85 | 18.5 | 31.7 | 52.8
| Standard Pressure (SP) (with vacuum) | 0.5 | 1.3 | 2.6 | 4.4 | 5.3 | 10.575 | 15.85 | 18.5 | 31.7 | 52.8
| Standard Pressure (SP) (no vacuum) | 0.5 | 1.3 | 2.6 | 4.4 | 5.3 | 10.575 | 15.85 | 18.5 | 31.7 | 52.8
| High Pressure (HP) | 0.5 | 1.3 | 2.6 | 4.4 | 5.3 | 10.575 | 15.85 | 18.5 | 31.7 | 52.8
5E1-S | Low Pressure (LP) (with vacuum) | 1.3 | 3.3 | 6.6 | 10.2 | 13.2 | 26.4 | 39.6 | 46.2 | 79.2 | 132
| Low Pressure (LP) (no vacuum) | 1.3 | 3.3 | 6.6 | 10.2 | 13.2 | 26.4 | 39.6 | 46.2 | 79.2 | 132
| Standard Pressure (SP) (with vacuum) | 1.3 | 3.3 | 6.6 | 10.2 | 13.2 | 26.4 | 39.6 | 46.2 | 79.2 | 132
| Standard Pressure (SP) (no vacuum) | 1.3 | 3.3 | 6.6 | 10.2 | 13.2 | 26.4 | 39.6 | 46.2 | 79.2 | 132
| High Pressure (HP) | 1.3 | 3.3 | 6.6 | 10.2 | 13.2 | 26.4 | 39.6 | 46.2 | 79.2 | 132
1E2-S | Low Pressure (LP) (with vacuum) | 2.6 | 6.6 | 13.2 | 20.1 | 26.4 | 52.8 | 79.2 | 92.4 | 158.4 | 264
| Low Pressure (LP) (no vacuum) | 2.6 | 6.6 | 13.2 | 20.1 | 26.4 | 52.8 | 79.2 | 92.4 | 158.4 | 264
| Standard Pressure (SP) (with vacuum) | 2.6 | 6.6 | 13.2 | 20.1 | 26.4 | 52.8 | 79.2 | 92.4 | 158.4 | 264
| Standard Pressure (SP) (no vacuum) | 2.6 | 6.6 | 13.2 | 20.1 | 26.4 | 52.8 | 79.2 | 92.4 | 158.4 | 264
| High Pressure (HP) | 2.6 | 6.6 | 13.2 | 20.1 | 26.4 | 52.8 | 79.2 | 92.4 | 158.4 | 264
2E2-S | Low Pressure (LP) (with vacuum) | 5.3 | 13.2 | 26.4 | 36.7 | 52.8 | 105.6 | 158.4 | 184.8 | 316.8 | 528.1
| Low Pressure (LP) (no vacuum) | 5.3 | 13.2 | 26.4 | 36.7 | 52.8 | 105.6 | 158.4 | 184.8 | 316.8 | 528.1
| Standard Pressure (SP) (with vacuum) | 5.3 | 13.2 | 26.4 | 36.7 | 52.8 | 105.6 | 158.4 | 184.8 | 316.8 | 528.1
| Standard Pressure (SP) (no vacuum) | 5.3 | 13.2 | 26.4 | 36.7 | 52.8 | 105.6 | 158.4 | 184.8 | 316.8 | 528.1
| High Pressure (HP) | 5.3 | 13.2 | 26.4 | 36.7 | 52.8 | 105.6 | 158.4 | 184.8 | 316.8 | 528.1
5E2-S | Low Pressure (LP) (with vacuum) | 13.2 | 33 | 66 | 88.2 | 132 | 264.025 | 396.05 | 462.1 | 792.1 | 1320.2
| Low Pressure (LP) (no vacuum) | 13.2 | 33 | 66 | 88.2 | 132 | 264.025 | 396.05 | 462.1 | 792.1 | 1320.2
| Standard Pressure (SP) (with vacuum) | 13.2 | 33 | 66 | 88.2 | 132 | 264.025 | 396.05 | 462.1 | 792.1 | 1320.2
| Standard Pressure (SP) (no vacuum) | 13.2 | 33 | 66 | 88.2 | 132 | 264.025 | 396.05 | 462.1 | 792.1 | 1320.2
| High Pressure (HP) | 13.2 | 33 | 66 | 88.2 | 132 | 264.025 | 396.05 | 462.1 | 792.1 | 1320.2
1E3-S | Low Pressure (LP) (with vacuum) | 26.4 | 66 | 132 | 170 | 264 | 528.05 | 792.1 | 924.1 | 1584.2 | 2640.3
| Low Pressure (LP) (no vacuum) | 26.4 | 66 | 132 | 170 | 264 | 528.05 | 792.1 | 924.1 | 1584.2 | 2640.3
| Standard Pressure (SP) (with vacuum) | 26.4 | 66 | 132 | 170 | 264 | 528.05 | 792.1 | 924.1 | 1584.2 | 2640.3
| Standard Pressure (SP) (no vacuum) | 26.4 | 66 | 132 | 170 | 264 | 528.05 | 792.1 | 924.1 | 1584.2 | 2640.3
| High Pressure (HP) | 26.4 | 66 | 132 | 170 | 264 | 528.05 | 792.1 | 924.1 | 1584.2 | 2640.3
2E3-S | Low Pressure (LP) (with vacuum) | 52.8 | 132 | 264 | 327.3 | 528.1 | 1056.15 | 1584.2 | 1848.2 | 3168.4 | 5280.6
| Low Pressure (LP) (no vacuum) | 52.8 | 132 | 264 | 327.3 | 528.1 | 1056.15 | 1584.2 | 1848.2 | 3168.4 | 5280.6
| Standard Pressure (SP) (with vacuum) | 52.8 | 132 | 264 | 327.3 | 528.1 | 1056.15 | 1584.2 | 1848.2 | 3168.4 | 5280.6
| Standard Pressure (SP) (no vacuum) | 52.8 | 132 | 264 | 327.3 | 528.1 | 1056.15 | 1584.2 | 1848.2 | 3168.4 | 5280.6
| High Pressure (HP) | 52.8 | 132 | 264 | 327.3 | 528.1 | 1056.15 | 1584.2 | 1848.2 | 3168.4 | 5280.6
5E3-S | Low Pressure (LP) (with vacuum) | 132 | 330 | 660.1 | 785.5 | 1320.2 | 2640.325 | 3960.45 | 4620.5 | 7920.9 | 13201.5
| Low Pressure (LP) (no vacuum) | 132 | 330 | 660.1 | 785.5 | 1320.2 | 2640.325 | 3960.45 | 4620.5 | 7920.9 | 13201.5
| Standard Pressure (SP) (with vacuum) | 132 | 330 | 660.1 | 785.5 | 1320.2 | 2640.325 | 3960.45 | 4620.5 | 7920.9 | 13201.5
| Standard Pressure (SP) (no vacuum) | 132 | 330 | 660.1 | 785.5 | 1320.2 | 2640.325 | 3960.45 | 4620.5 | 7920.9 | 13201.5
| High Pressure (HP) | 132 | 330 | 660.1 | 785.5 | 1320.2 | 2640.325 | 3960.45 | 4620.5 | 7920.9 | 13201.5
1E4-S | Low Pressure (LP) (with vacuum) | 264 | 660.1 | 1320.2 | 1517.2 | 2640.3 | 5280.625 | 7920.95 | 9241.1 | 15841.8 | 26403
| Low Pressure (LP) (no vacuum) | 264 | 660.1 | 1320.2 | 1517.2 | 2640.3 | 5280.625 | 7920.95 | 9241.1 | 15841.8 | 26403
| Standard Pressure (SP) (with vacuum) | 264 | 660.1 | 1320.2 | 1517.2 | 2640.3 | 5280.625 | 7920.95 | 9241.1 | 15841.8 | 26403
| Standard Pressure (SP) (no vacuum) | 264 | 660.1 | 1320.2 | 1517.2 | 2640.3 | 5280.625 | 7920.95 | 9241.1 | 15841.8 | 26403
| High Pressure (HP) | 264 | 660.1 | 1320.2 | 1517.2 | 2640.3 | 5280.625 | 7920.95 | 9241.1 | 15841.8 | 26403
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.
Propane Ratio = 0.789 | C3H8 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.4 | 0.8 | 1.3 | 1.6 | 3.15 | 4.7 | 5.5 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.2 | 0.4 | 0.8 | 1.3 | 1.6 | 3.15 | 4.7 | 5.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.4 | 0.8 | 1.3 | 1.6 | 3.15 | 4.7 | 5.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.4 | 0.8 | 1.3 | 1.6 | 3.15 | 4.7 | 5.5 |  | n/a[4]
| High Pressure (HP) | 0.2 | 0.4 | 0.8 | 1.3 | 1.6 | 3.15 | 4.7 | 5.5 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.3 | 0.8 | 1.6 | 2.3 | 3.2 | 6.325 | 9.45 | 11 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.3 | 0.8 | 1.6 | 2.3 | 3.2 | 6.325 | 9.45 | 11 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.3 | 0.8 | 1.6 | 2.3 | 3.2 | 6.325 | 9.45 | 11 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.3 | 0.8 | 1.6 | 2.3 | 3.2 | 6.325 | 9.45 | 11 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.3 | 0.8 | 1.6 | 2.3 | 3.2 | 6.325 | 9.45 | 11 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.8 | 2 | 3.9 | 5.5 | 7.9 | 15.775 | 23.65 | 27.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.8 | 2 | 3.9 | 5.5 | 7.9 | 15.775 | 23.65 | 27.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.8 | 2 | 3.9 | 5.5 | 7.9 | 15.775 | 23.65 | 27.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.8 | 2 | 3.9 | 5.5 | 7.9 | 15.775 | 23.65 | 27.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.8 | 2 | 3.9 | 5.5 | 7.9 | 15.775 | 23.65 | 27.6 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 1.6 | 3.9 | 7.9 | 10.5 | 15.8 | 31.55 | 47.3 | 55.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.6 | 3.9 | 7.9 | 10.5 | 15.8 | 31.55 | 47.3 | 55.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.6 | 3.9 | 7.9 | 10.5 | 15.8 | 31.55 | 47.3 | 55.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.6 | 3.9 | 7.9 | 10.5 | 15.8 | 31.55 | 47.3 | 55.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.6 | 3.9 | 7.9 | 10.5 | 15.8 | 31.55 | 47.3 | 55.2 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 3.2 | 7.9 | 15.8 | 20.8 | 31.5 | 63.075 | 94.65 | 110.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 3.2 | 7.9 | 15.8 | 20.8 | 31.5 | 63.075 | 94.65 | 110.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 3.2 | 7.9 | 15.8 | 20.8 | 31.5 | 63.075 | 94.65 | 110.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 3.2 | 7.9 | 15.8 | 20.8 | 31.5 | 63.075 | 94.65 | 110.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 3.2 | 7.9 | 15.8 | 20.8 | 31.5 | 63.075 | 94.65 | 110.4 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 7.9 | 19.7 | 39.4 | 48.8 | 78.9 | 157.775 | 236.65 | 276.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 7.9 | 19.7 | 39.4 | 48.8 | 78.9 | 157.775 | 236.65 | 276.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 7.9 | 19.7 | 39.4 | 48.8 | 78.9 | 157.775 | 236.65 | 276.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 7.9 | 19.7 | 39.4 | 48.8 | 78.9 | 157.775 | 236.65 | 276.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 7.9 | 19.7 | 39.4 | 48.8 | 78.9 | 157.775 | 236.65 | 276.1 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 15.8 | 39.4 | 78.9 | 97.5 | 157.7 | 315.475 | 473.25 | 552.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 15.8 | 39.4 | 78.9 | 97.5 | 157.7 | 315.475 | 473.25 | 552.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 15.8 | 39.4 | 78.9 | 97.5 | 157.7 | 315.475 | 473.25 | 552.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 15.8 | 39.4 | 78.9 | 97.5 | 157.7 | 315.475 | 473.25 | 552.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 15.8 | 39.4 | 78.9 | 97.5 | 157.7 | 315.475 | 473.25 | 552.1 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 31.5 | 78.9 | 157.7 | 181.4 | 315.5 | 630.975 | 946.45 | 1104.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 31.5 | 78.9 | 157.7 | 181.4 | 315.5 | 630.975 | 946.45 | 1104.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 31.5 | 78.9 | 157.7 | 181.4 | 315.5 | 630.975 | 946.45 | 1104.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 31.5 | 78.9 | 157.7 | 181.4 | 315.5 | 630.975 | 946.45 | 1104.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 31.5 | 78.9 | 157.7 | 181.4 | 315.5 | 630.975 | 946.45 | 1104.2 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 78.9 | 197.2 | 394.4 | 453.2 | 788.7 | 1577.425 | 2366.15 | 2760.5 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 78.9 | 197.2 | 394.4 | 453.2 | 788.7 | 1577.425 | 2366.15 | 2760.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 78.9 | 197.2 | 394.4 | 453.2 | 788.7 | 1577.425 | 2366.15 | 2760.5 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 78.9 | 197.2 | 394.4 | 453.2 | 788.7 | 1577.425 | 2366.15 | 2760.5 | n/a[4] | n/a[4]
| High Pressure (HP) | 78.9 | 197.2 | 394.4 | 453.2 | 788.7 | 1577.425 | 2366.15 | 2760.5 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 157.7 | 394.4 | 788.7 | 906.5 | 1577.4 | 3154.85 | 4732.3 | 5521 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 157.7 | 394.4 | 788.7 | 906.5 | 1577.4 | 3154.85 | 4732.3 | 5521 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 157.7 | 394.4 | 788.7 | 906.5 | 1577.4 | 3154.85 | 4732.3 | 5521 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 157.7 | 394.4 | 788.7 | 906.5 | 1577.4 | 3154.85 | 4732.3 | 5521 | n/a[4] | n/a[4]
| High Pressure (HP) | 157.7 | 394.4 | 788.7 | 906.5 | 1577.4 | 3154.85 | 4732.3 | 5521 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Carbon Tetrafluoride Ratio = 0.563 | CF4 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.3 | 0.6 | 0.9 | 1.1 | 2.225 | 3.35 | 3.9 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.3 | 0.6 | 0.9 | 1.1 | 2.225 | 3.35 | 3.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.3 | 0.6 | 0.9 | 1.1 | 2.225 | 3.35 | 3.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.3 | 0.6 | 0.9 | 1.1 | 2.225 | 3.35 | 3.9 |  | n/a[4]
| High Pressure (HP) | 0.1 | 0.3 | 0.6 | 0.9 | 1.1 | 2.225 | 3.35 | 3.9 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.6 | 1.1 | 1.8 | 2.3 | 4.525 | 6.75 | 7.9 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.2 | 0.6 | 1.1 | 1.8 | 2.3 | 4.525 | 6.75 | 7.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.6 | 1.1 | 1.8 | 2.3 | 4.525 | 6.75 | 7.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.6 | 1.1 | 1.8 | 2.3 | 4.525 | 6.75 | 7.9 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.2 | 0.6 | 1.1 | 1.8 | 2.3 | 4.525 | 6.75 | 7.9 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.6 | 1.4 | 2.8 | 4.1 | 5.6 | 11.25 | 16.9 | 19.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.6 | 1.4 | 2.8 | 4.1 | 5.6 | 11.25 | 16.9 | 19.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.6 | 1.4 | 2.8 | 4.1 | 5.6 | 11.25 | 16.9 | 19.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.6 | 1.4 | 2.8 | 4.1 | 5.6 | 11.25 | 16.9 | 19.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.6 | 1.4 | 2.8 | 4.1 | 5.6 | 11.25 | 16.9 | 19.7 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 1.1 | 2.8 | 5.6 | 7.9 | 11.3 | 22.525 | 33.75 | 39.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.1 | 2.8 | 5.6 | 7.9 | 11.3 | 22.525 | 33.75 | 39.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.1 | 2.8 | 5.6 | 7.9 | 11.3 | 22.525 | 33.75 | 39.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.1 | 2.8 | 5.6 | 7.9 | 11.3 | 22.525 | 33.75 | 39.4 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.1 | 2.8 | 5.6 | 7.9 | 11.3 | 22.525 | 33.75 | 39.4 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 2.3 | 5.6 | 11.3 | 15.7 | 22.5 | 45.025 | 67.55 | 78.8 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 2.3 | 5.6 | 11.3 | 15.7 | 22.5 | 45.025 | 67.55 | 78.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 2.3 | 5.6 | 11.3 | 15.7 | 22.5 | 45.025 | 67.55 | 78.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 2.3 | 5.6 | 11.3 | 15.7 | 22.5 | 45.025 | 67.55 | 78.8 | n/a[4] | n/a[4]
| High Pressure (HP) | 2.3 | 5.6 | 11.3 | 15.7 | 22.5 | 45.025 | 67.55 | 78.8 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 5.6 | 14.1 | 28.1 | 36.3 | 56.3 | 112.575 | 168.85 | 197 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 5.6 | 14.1 | 28.1 | 36.3 | 56.3 | 112.575 | 168.85 | 197 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 5.6 | 14.1 | 28.1 | 36.3 | 56.3 | 112.575 | 168.85 | 197 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 5.6 | 14.1 | 28.1 | 36.3 | 56.3 | 112.575 | 168.85 | 197 | n/a[4] | n/a[4]
| High Pressure (HP) | 5.6 | 14.1 | 28.1 | 36.3 | 56.3 | 112.575 | 168.85 | 197 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 11.3 | 28.1 | 56.3 | 69.9 | 112.6 | 225.15 | 337.7 | 394 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 11.3 | 28.1 | 56.3 | 69.9 | 112.6 | 225.15 | 337.7 | 394 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 11.3 | 28.1 | 56.3 | 69.9 | 112.6 | 225.15 | 337.7 | 394 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 11.3 | 28.1 | 56.3 | 69.9 | 112.6 | 225.15 | 337.7 | 394 | n/a[4] | n/a[4]
| High Pressure (HP) | 11.3 | 28.1 | 56.3 | 69.9 | 112.6 | 225.15 | 337.7 | 394 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 22.5 | 56.3 | 112.6 | 134.1 | 225.1 | 450.275 | 675.45 | 788 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 22.5 | 56.3 | 112.6 | 134.1 | 225.1 | 450.275 | 675.45 | 788 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 22.5 | 56.3 | 112.6 | 134.1 | 225.1 | 450.275 | 675.45 | 788 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 22.5 | 56.3 | 112.6 | 134.1 | 225.1 | 450.275 | 675.45 | 788 | n/a[4] | n/a[4]
| High Pressure (HP) | 22.5 | 56.3 | 112.6 | 134.1 | 225.1 | 450.275 | 675.45 | 788 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 56.3 | 140.7 | 281.4 | 323.4 | 562.9 | 1125.75 | 1688.6 | 1970 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 56.3 | 140.7 | 281.4 | 323.4 | 562.9 | 1125.75 | 1688.6 | 1970 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 56.3 | 140.7 | 281.4 | 323.4 | 562.9 | 1125.75 | 1688.6 | 1970 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 56.3 | 140.7 | 281.4 | 323.4 | 562.9 | 1125.75 | 1688.6 | 1970 | n/a[4] | n/a[4]
| High Pressure (HP) | 56.3 | 140.7 | 281.4 | 323.4 | 562.9 | 1125.75 | 1688.6 | 1970 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 112.6 | 281.4 | 562.9 | 646.9 | 1125.7 | 2251.45 | 3377.2 | 3940.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 112.6 | 281.4 | 562.9 | 646.9 | 1125.7 | 2251.45 | 3377.2 | 3940.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 112.6 | 281.4 | 562.9 | 646.9 | 1125.7 | 2251.45 | 3377.2 | 3940.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 112.6 | 281.4 | 562.9 | 646.9 | 1125.7 | 2251.45 | 3377.2 | 3940.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 112.6 | 281.4 | 562.9 | 646.9 | 1125.7 | 2251.45 | 3377.2 | 3940.1 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Hexafluoroethane Ratio = 0.447 | C2F6 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.775 | 2.65 | 3.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.775 | 2.65 | 3.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.775 | 2.65 | 3.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.775 | 2.65 | 3.1 |  | n/a[4]
| High Pressure (HP) | 0.1 | 0.2 | 0.4 | 0.8 | 0.9 | 1.775 | 2.65 | 3.1 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.4 | 0.9 | 1.3 | 1.8 | 3.6 | 5.4 | 6.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.2 | 0.4 | 0.9 | 1.3 | 1.8 | 3.6 | 5.4 | 6.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.4 | 0.9 | 1.3 | 1.8 | 3.6 | 5.4 | 6.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.4 | 0.9 | 1.3 | 1.8 | 3.6 | 5.4 | 6.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.2 | 0.4 | 0.9 | 1.3 | 1.8 | 3.6 | 5.4 | 6.3 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 1.1 | 2.2 | 3.2 | 4.5 | 8.975 | 13.45 | 15.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.4 | 1.1 | 2.2 | 3.2 | 4.5 | 8.975 | 13.45 | 15.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.4 | 1.1 | 2.2 | 3.2 | 4.5 | 8.975 | 13.45 | 15.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.4 | 1.1 | 2.2 | 3.2 | 4.5 | 8.975 | 13.45 | 15.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.4 | 1.1 | 2.2 | 3.2 | 4.5 | 8.975 | 13.45 | 15.7 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 0.9 | 2.2 | 4.5 | 5.9 | 8.9 | 17.875 | 26.85 | 31.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.9 | 2.2 | 4.5 | 5.9 | 8.9 | 17.875 | 26.85 | 31.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.9 | 2.2 | 4.5 | 5.9 | 8.9 | 17.875 | 26.85 | 31.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.9 | 2.2 | 4.5 | 5.9 | 8.9 | 17.875 | 26.85 | 31.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.9 | 2.2 | 4.5 | 5.9 | 8.9 | 17.875 | 26.85 | 31.3 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 1.8 | 4.5 | 8.9 | 11.8 | 17.9 | 35.775 | 53.65 | 62.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.8 | 4.5 | 8.9 | 11.8 | 17.9 | 35.775 | 53.65 | 62.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.8 | 4.5 | 8.9 | 11.8 | 17.9 | 35.775 | 53.65 | 62.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.8 | 4.5 | 8.9 | 11.8 | 17.9 | 35.775 | 53.65 | 62.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.8 | 4.5 | 8.9 | 11.8 | 17.9 | 35.775 | 53.65 | 62.6 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 4.5 | 11.2 | 22.4 | 27.6 | 44.7 | 89.45 | 134.2 | 156.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 4.5 | 11.2 | 22.4 | 27.6 | 44.7 | 89.45 | 134.2 | 156.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 4.5 | 11.2 | 22.4 | 27.6 | 44.7 | 89.45 | 134.2 | 156.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 4.5 | 11.2 | 22.4 | 27.6 | 44.7 | 89.45 | 134.2 | 156.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 4.5 | 11.2 | 22.4 | 27.6 | 44.7 | 89.45 | 134.2 | 156.6 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 8.9 | 22.4 | 44.7 | 55.3 | 89.5 | 178.95 | 268.4 | 313.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 8.9 | 22.4 | 44.7 | 55.3 | 89.5 | 178.95 | 268.4 | 313.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 8.9 | 22.4 | 44.7 | 55.3 | 89.5 | 178.95 | 268.4 | 313.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 8.9 | 22.4 | 44.7 | 55.3 | 89.5 | 178.95 | 268.4 | 313.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 8.9 | 22.4 | 44.7 | 55.3 | 89.5 | 178.95 | 268.4 | 313.1 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 17.9 | 44.7 | 89.5 | 102.9 | 178.9 | 357.85 | 536.8 | 626.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 17.9 | 44.7 | 89.5 | 102.9 | 178.9 | 357.85 | 536.8 | 626.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 17.9 | 44.7 | 89.5 | 102.9 | 178.9 | 357.85 | 536.8 | 626.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 17.9 | 44.7 | 89.5 | 102.9 | 178.9 | 357.85 | 536.8 | 626.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 17.9 | 44.7 | 89.5 | 102.9 | 178.9 | 357.85 | 536.8 | 626.3 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 44.7 | 111.8 | 223.7 | 257.1 | 447.3 | 894.65 | 1342 | 1565.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 44.7 | 111.8 | 223.7 | 257.1 | 447.3 | 894.65 | 1342 | 1565.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 44.7 | 111.8 | 223.7 | 257.1 | 447.3 | 894.65 | 1342 | 1565.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 44.7 | 111.8 | 223.7 | 257.1 | 447.3 | 894.65 | 1342 | 1565.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 44.7 | 111.8 | 223.7 | 257.1 | 447.3 | 894.65 | 1342 | 1565.7 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 89.5 | 223.7 | 447.3 | 514.1 | 894.7 | 1789.35 | 2684 | 3131.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 89.5 | 223.7 | 447.3 | 514.1 | 894.7 | 1789.35 | 2684 | 3131.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 89.5 | 223.7 | 447.3 | 514.1 | 894.7 | 1789.35 | 2684 | 3131.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 89.5 | 223.7 | 447.3 | 514.1 | 894.7 | 1789.35 | 2684 | 3131.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 89.5 | 223.7 | 447.3 | 514.1 | 894.7 | 1789.35 | 2684 | 3131.3 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Trifluoromethane Ratio = 0.629 | CHF3 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.3 | 0.6 | 1 | 1.3 | 2.525 | 3.75 | 4.4 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.3 | 0.6 | 1 | 1.3 | 2.525 | 3.75 | 4.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.3 | 0.6 | 1 | 1.3 | 2.525 | 3.75 | 4.4 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.3 | 0.6 | 1 | 1.3 | 2.525 | 3.75 | 4.4 |  | n/a[4]
| High Pressure (HP) | 0.1 | 0.3 | 0.6 | 1 | 1.3 | 2.525 | 3.75 | 4.4 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.3 | 0.6 | 1.3 | 2 | 2.5 | 5.025 | 7.55 | 8.8 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.3 | 0.6 | 1.3 | 2 | 2.5 | 5.025 | 7.55 | 8.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.3 | 0.6 | 1.3 | 2 | 2.5 | 5.025 | 7.55 | 8.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.3 | 0.6 | 1.3 | 2 | 2.5 | 5.025 | 7.55 | 8.8 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.3 | 0.6 | 1.3 | 2 | 2.5 | 5.025 | 7.55 | 8.8 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.6 | 1.6 | 3.1 | 4.4 | 6.3 | 12.575 | 18.85 | 22 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.6 | 1.6 | 3.1 | 4.4 | 6.3 | 12.575 | 18.85 | 22 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.6 | 1.6 | 3.1 | 4.4 | 6.3 | 12.575 | 18.85 | 22 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.6 | 1.6 | 3.1 | 4.4 | 6.3 | 12.575 | 18.85 | 22 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.6 | 1.6 | 3.1 | 4.4 | 6.3 | 12.575 | 18.85 | 22 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 1.3 | 3.1 | 6.3 | 8.8 | 12.6 | 25.2 | 37.8 | 44.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.3 | 3.1 | 6.3 | 8.8 | 12.6 | 25.2 | 37.8 | 44.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.3 | 3.1 | 6.3 | 8.8 | 12.6 | 25.2 | 37.8 | 44.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.3 | 3.1 | 6.3 | 8.8 | 12.6 | 25.2 | 37.8 | 44.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.3 | 3.1 | 6.3 | 8.8 | 12.6 | 25.2 | 37.8 | 44.1 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 2.5 | 6.3 | 12.6 | 17.2 | 25.2 | 50.35 | 75.5 | 88.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 2.5 | 6.3 | 12.6 | 17.2 | 25.2 | 50.35 | 75.5 | 88.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 2.5 | 6.3 | 12.6 | 17.2 | 25.2 | 50.35 | 75.5 | 88.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 2.5 | 6.3 | 12.6 | 17.2 | 25.2 | 50.35 | 75.5 | 88.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 2.5 | 6.3 | 12.6 | 17.2 | 25.2 | 50.35 | 75.5 | 88.1 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 6.3 | 15.7 | 31.5 | 40.6 | 62.9 | 125.85 | 188.8 | 220.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 6.3 | 15.7 | 31.5 | 40.6 | 62.9 | 125.85 | 188.8 | 220.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 6.3 | 15.7 | 31.5 | 40.6 | 62.9 | 125.85 | 188.8 | 220.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 6.3 | 15.7 | 31.5 | 40.6 | 62.9 | 125.85 | 188.8 | 220.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 6.3 | 15.7 | 31.5 | 40.6 | 62.9 | 125.85 | 188.8 | 220.3 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 12.6 | 31.5 | 62.9 | 77.9 | 125.9 | 251.775 | 377.65 | 440.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 12.6 | 31.5 | 62.9 | 77.9 | 125.9 | 251.775 | 377.65 | 440.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 12.6 | 31.5 | 62.9 | 77.9 | 125.9 | 251.775 | 377.65 | 440.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 12.6 | 31.5 | 62.9 | 77.9 | 125.9 | 251.775 | 377.65 | 440.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 12.6 | 31.5 | 62.9 | 77.9 | 125.9 | 251.775 | 377.65 | 440.6 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 25.2 | 62.9 | 125.9 | 149.9 | 251.8 | 503.525 | 755.25 | 881.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 25.2 | 62.9 | 125.9 | 149.9 | 251.8 | 503.525 | 755.25 | 881.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 25.2 | 62.9 | 125.9 | 149.9 | 251.8 | 503.525 | 755.25 | 881.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 25.2 | 62.9 | 125.9 | 149.9 | 251.8 | 503.525 | 755.25 | 881.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 25.2 | 62.9 | 125.9 | 149.9 | 251.8 | 503.525 | 755.25 | 881.1 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 62.9 | 157.3 | 314.7 | 361.7 | 629.4 | 1258.75 | 1888.1 | 2202.8 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 62.9 | 157.3 | 314.7 | 361.7 | 629.4 | 1258.75 | 1888.1 | 2202.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 62.9 | 157.3 | 314.7 | 361.7 | 629.4 | 1258.75 | 1888.1 | 2202.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 62.9 | 157.3 | 314.7 | 361.7 | 629.4 | 1258.75 | 1888.1 | 2202.8 | n/a[4] | n/a[4]
| High Pressure (HP) | 62.9 | 157.3 | 314.7 | 361.7 | 629.4 | 1258.75 | 1888.1 | 2202.8 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 125.9 | 314.7 | 629.4 | 723.3 | 1258.8 | 2517.525 | 3776.25 | 4405.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 125.9 | 314.7 | 629.4 | 723.3 | 1258.8 | 2517.525 | 3776.25 | 4405.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 125.9 | 314.7 | 629.4 | 723.3 | 1258.8 | 2517.525 | 3776.25 | 4405.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 125.9 | 314.7 | 629.4 | 723.3 | 1258.8 | 2517.525 | 3776.25 | 4405.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 125.9 | 314.7 | 629.4 | 723.3 | 1258.8 | 2517.525 | 3776.25 | 4405.6 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Carbon Monoxide Ratio = 1.000 | CO | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.5 | 1 | 2 | 2 | 4 | 6 | 7 | 12 | 20
| Low Pressure (LP) (no vacuum) | 0.2 | 0.5 | 1 | 2 | 2 | 4 | 6 | 7 | 12 | 20
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.5 | 1 | 2 | 2 | 4 | 6 | 7 | 12 | 20
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.5 | 1 | 2 | 2 | 4 | 6 | 7 | 12 | 20
| High Pressure (HP) | 0.2 | 0.5 | 1 | 2 | 2 | 4 | 6 | 7 | 12 | 20
2E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
| Low Pressure (LP) (no vacuum) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
| Standard Pressure (SP) (with vacuum) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
| Standard Pressure (SP) (no vacuum) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
| High Pressure (HP) | 0.4 | 1 | 2 | 3.5 | 4 | 8 | 12 | 14 | 24 | 40
5E1-S | Low Pressure (LP) (with vacuum) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
| Low Pressure (LP) (no vacuum) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
| Standard Pressure (SP) (with vacuum) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
| Standard Pressure (SP) (no vacuum) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
| High Pressure (HP) | 1 | 2.5 | 5 | 7.7 | 10 | 20 | 30 | 35 | 60 | 100
1E2-S | Low Pressure (LP) (with vacuum) | 2 | 5 | 10 | 15.4 | 20 | 40 | 60 | 70 | 120 | 200
| Low Pressure (LP) (no vacuum) | 2 | 5 | 10 | 15.4 | 20 | 40 | 60 | 70 | 120 | 200
| Standard Pressure (SP) (with vacuum) | 2 | 5 | 10 | 15.4 | 20 | 40 | 60 | 70 | 120 | 200
| Standard Pressure (SP) (no vacuum) | 2 | 5 | 10 | 15.4 | 20 | 40 | 60 | 70 | 120 | 200
| High Pressure (HP) | 2 | 5 | 10 | 15.4 | 20 | 40 | 60 | 70 | 120 | 200
2E2-S | Low Pressure (LP) (with vacuum) | 4 | 10 | 20 | 27.8 | 40 | 80 | 120 | 140 | 240 | 400
| Low Pressure (LP) (no vacuum) | 4 | 10 | 20 | 27.8 | 40 | 80 | 120 | 140 | 240 | 400
| Standard Pressure (SP) (with vacuum) | 4 | 10 | 20 | 27.8 | 40 | 80 | 120 | 140 | 240 | 400
| Standard Pressure (SP) (no vacuum) | 4 | 10 | 20 | 27.8 | 40 | 80 | 120 | 140 | 240 | 400
| High Pressure (HP) | 4 | 10 | 20 | 27.8 | 40 | 80 | 120 | 140 | 240 | 400
5E2-S | Low Pressure (LP) (with vacuum) | 10 | 25 | 50 | 68.3 | 100 | 200 | 300 | 350 | 600 | 1000
| Low Pressure (LP) (no vacuum) | 10 | 25 | 50 | 68.3 | 100 | 200 | 300 | 350 | 600 | 1000
| Standard Pressure (SP) (with vacuum) | 10 | 25 | 50 | 68.3 | 100 | 200 | 300 | 350 | 600 | 1000
| Standard Pressure (SP) (no vacuum) | 10 | 25 | 50 | 68.3 | 100 | 200 | 300 | 350 | 600 | 1000
| High Pressure (HP) | 10 | 25 | 50 | 68.3 | 100 | 200 | 300 | 350 | 600 | 1000
1E3-S | Low Pressure (LP) (with vacuum) | 20 | 50 | 100 | 128.7 | 200 | 400 | 600 | 700 | 1200 | 1999.9
| Low Pressure (LP) (no vacuum) | 20 | 50 | 100 | 128.7 | 200 | 400 | 600 | 700 | 1200 | 1999.9
| Standard Pressure (SP) (with vacuum) | 20 | 50 | 100 | 128.7 | 200 | 400 | 600 | 700 | 1200 | 1999.9
| Standard Pressure (SP) (no vacuum) | 20 | 50 | 100 | 128.7 | 200 | 400 | 600 | 700 | 1200 | 1999.9
| High Pressure (HP) | 20 | 50 | 100 | 128.7 | 200 | 400 | 600 | 700 | 1200 | 1999.9
2E3-S | Low Pressure (LP) (with vacuum) | 40 | 100 | 200 | 247.9 | 400 | 799.975 | 1199.95 | 1399.9 | 2399.9 | 3999.9
| Low Pressure (LP) (no vacuum) | 40 | 100 | 200 | 247.9 | 400 | 799.975 | 1199.95 | 1399.9 | 2399.9 | 3999.9
| Standard Pressure (SP) (with vacuum) | 40 | 100 | 200 | 247.9 | 400 | 799.975 | 1199.95 | 1399.9 | 2399.9 | 3999.9
| Standard Pressure (SP) (no vacuum) | 40 | 100 | 200 | 247.9 | 400 | 799.975 | 1199.95 | 1399.9 | 2399.9 | 3999.9
| High Pressure (HP) | 40 | 100 | 200 | 247.9 | 400 | 799.975 | 1199.95 | 1399.9 | 2399.9 | 3999.9
5E3-S | Low Pressure (LP) (with vacuum) | 100 | 250 | 500 | 595 | 1000 | 1999.95 | 2999.9 | 3499.9 | 5999.8 | 9999.6
| Low Pressure (LP) (no vacuum) | 100 | 250 | 500 | 595 | 1000 | 1999.95 | 2999.9 | 3499.9 | 5999.8 | 9999.6
| Standard Pressure (SP) (with vacuum) | 100 | 250 | 500 | 595 | 1000 | 1999.95 | 2999.9 | 3499.9 | 5999.8 | 9999.6
| Standard Pressure (SP) (no vacuum) | 100 | 250 | 500 | 595 | 1000 | 1999.95 | 2999.9 | 3499.9 | 5999.8 | 9999.6
| High Pressure (HP) | 100 | 250 | 500 | 595 | 1000 | 1999.95 | 2999.9 | 3499.9 | 5999.8 | 9999.6
1E4-S | Low Pressure (LP) (with vacuum) | 200 | 500 | 1000 | 1190 | 1999.9 | 3999.825 | 5999.75 | 6999.7 | 11999.6 | 19999.3
| Low Pressure (LP) (no vacuum) | 200 | 500 | 1000 | 1190 | 1999.9 | 3999.825 | 5999.75 | 6999.7 | 11999.6 | 19999.3
| Standard Pressure (SP) (with vacuum) | 200 | 500 | 1000 | 1190 | 1999.9 | 3999.825 | 5999.75 | 6999.7 | 11999.6 | 19999.3
| Standard Pressure (SP) (no vacuum) | 200 | 500 | 1000 | 1190 | 1999.9 | 3999.825 | 5999.75 | 6999.7 | 11999.6 | 19999.3
| High Pressure (HP) | 200 | 500 | 1000 | 1190 | 1999.9 | 3999.825 | 5999.75 | 6999.7 | 11999.6 | 19999.3
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.
Nitrous Oxide Ratio = 0.795 | N2O | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 |  | n/a[4]
| High Pressure (HP) | 0.2 | 0.4 | 0.8 | 1.4 | 1.6 | 3.2 | 4.8 | 5.6 | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.3 | 0.8 | 1.6 | 2.5 | 3.2 | 6.375 | 9.55 | 11.1 | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[4] | n/a[4]
| High Pressure (HP) | 0.8 | 2 | 4 | 6.2 | 8 | 15.925 | 23.85 | 27.8 | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[4] | n/a[4]
| High Pressure (HP) | 1.6 | 4 | 8 | 11.1 | 15.9 | 31.825 | 47.75 | 55.7 | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 3.2 | 8 | 15.9 | 22.1 | 31.8 | 63.6 | 95.4 | 111.3 | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.025 | 238.55 | 278.3 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.025 | 238.55 | 278.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.025 | 238.55 | 278.3 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.025 | 238.55 | 278.3 | n/a[4] | n/a[4]
| High Pressure (HP) | 8 | 19.9 | 39.8 | 51.2 | 79.5 | 159.025 | 238.55 | 278.3 | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 15.9 | 39.8 | 79.5 | 102.4 | 159 | 318.05 | 477.1 | 556.6 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 15.9 | 39.8 | 79.5 | 102.4 | 159 | 318.05 | 477.1 | 556.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 15.9 | 39.8 | 79.5 | 102.4 | 159 | 318.05 | 477.1 | 556.6 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 15.9 | 39.8 | 79.5 | 102.4 | 159 | 318.05 | 477.1 | 556.6 | n/a[4] | n/a[4]
| High Pressure (HP) | 15.9 | 39.8 | 79.5 | 102.4 | 159 | 318.05 | 477.1 | 556.6 | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 31.8 | 79.5 | 159 | 189.3 | 318 | 636.075 | 954.15 | 1113.2 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 31.8 | 79.5 | 159 | 189.3 | 318 | 636.075 | 954.15 | 1113.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 31.8 | 79.5 | 159 | 189.3 | 318 | 636.075 | 954.15 | 1113.2 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 31.8 | 79.5 | 159 | 189.3 | 318 | 636.075 | 954.15 | 1113.2 | n/a[4] | n/a[4]
| High Pressure (HP) | 31.8 | 79.5 | 159 | 189.3 | 318 | 636.075 | 954.15 | 1113.2 | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 79.5 | 198.8 | 397.6 | 473.1 | 795.1 | 1590.225 | 2385.35 | 2782.9 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 79.5 | 198.8 | 397.6 | 473.1 | 795.1 | 1590.225 | 2385.35 | 2782.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 79.5 | 198.8 | 397.6 | 473.1 | 795.1 | 1590.225 | 2385.35 | 2782.9 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 79.5 | 198.8 | 397.6 | 473.1 | 795.1 | 1590.225 | 2385.35 | 2782.9 | n/a[4] | n/a[4]
| High Pressure (HP) | 79.5 | 198.8 | 397.6 | 473.1 | 795.1 | 1590.225 | 2385.35 | 2782.9 | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 159 | 397.6 | 795.1 | 913.7 | 1590.2 | 3180.45 | 4770.7 | 5565.8 | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 159 | 397.6 | 795.1 | 913.7 | 1590.2 | 3180.45 | 4770.7 | 5565.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 159 | 397.6 | 795.1 | 913.7 | 1590.2 | 3180.45 | 4770.7 | 5565.8 | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 159 | 397.6 | 795.1 | 913.7 | 1590.2 | 3180.45 | 4770.7 | 5565.8 | n/a[4] | n/a[4]
| High Pressure (HP) | 159 | 397.6 | 795.1 | 913.7 | 1590.2 | 3180.45 | 4770.7 | 5565.8 | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Octafluorocyclobutane1 Ratio = 0.367 | C4F8 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.2 | 0.4 | 0.6 | 0.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.2 | 0.4 | 0.6 | 0.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.2 | 0.4 | 0.6 | 0.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.2 | 0.4 | 0.6 | 0.7 | n/a[8] | n/a[8] |  |  | n/a[4]
| High Pressure (HP) | 0.1 | 0.2 | 0.4 | 0.6 | 0.7 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
2E1-S | Low Pressure (LP) (with vacuum) | 0.1 | 0.4 | 0.7 | 1.1 | 1.5 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.1 | 0.4 | 0.7 | 1.1 | 1.5 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.1 | 0.4 | 0.7 | 1.1 | 1.5 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.1 | 0.4 | 0.7 | 1.1 | 1.5 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 0.1 | 0.4 | 0.7 | 1.1 | 1.5 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
5E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 0.9 | 1.8 | 2.4 | 3.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.4 | 0.9 | 1.8 | 2.4 | 3.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.4 | 0.9 | 1.8 | 2.4 | 3.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.4 | 0.9 | 1.8 | 2.4 | 3.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 0.4 | 0.9 | 1.8 | 2.4 | 3.7 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
1E2-S | Low Pressure (LP) (with vacuum) | 0.7 | 1.8 | 3.7 | 4.8 | 7.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 0.7 | 1.8 | 3.7 | 4.8 | 7.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 0.7 | 1.8 | 3.7 | 4.8 | 7.3 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 0.7 | 1.8 | 3.7 | 4.8 | 7.3 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 0.7 | 1.8 | 3.7 | 4.8 | 7.3 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
2E2-S | Low Pressure (LP) (with vacuum) | 1.5 | 3.7 | 7.3 | 9.2 | 14.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 1.5 | 3.7 | 7.3 | 9.2 | 14.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 1.5 | 3.7 | 7.3 | 9.2 | 14.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 1.5 | 3.7 | 7.3 | 9.2 | 14.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 1.5 | 3.7 | 7.3 | 9.2 | 14.7 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
5E2-S | Low Pressure (LP) (with vacuum) | 3.7 | 9.2 | 18.3 | 22.7 | 36.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 3.7 | 9.2 | 18.3 | 22.7 | 36.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 3.7 | 9.2 | 18.3 | 22.7 | 36.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 3.7 | 9.2 | 18.3 | 22.7 | 36.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 3.7 | 9.2 | 18.3 | 22.7 | 36.7 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
1E3-S | Low Pressure (LP) (with vacuum) | 7.3 | 18.3 | 36.7 | 43.8 | 73.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 7.3 | 18.3 | 36.7 | 43.8 | 73.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 7.3 | 18.3 | 36.7 | 43.8 | 73.3 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 7.3 | 18.3 | 36.7 | 43.8 | 73.3 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 7.3 | 18.3 | 36.7 | 43.8 | 73.3 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
2E3-S | Low Pressure (LP) (with vacuum) | 14.7 | 36.7 | 73.3 | 84.3 | 146.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 14.7 | 36.7 | 73.3 | 84.3 | 146.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 14.7 | 36.7 | 73.3 | 84.3 | 146.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 14.7 | 36.7 | 73.3 | 84.3 | 146.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 14.7 | 36.7 | 73.3 | 84.3 | 146.7 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
5E3-S | Low Pressure (LP) (with vacuum) | 36.7 | 91.7 | 183.3 | 210.7 | 366.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 36.7 | 91.7 | 183.3 | 210.7 | 366.7 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 36.7 | 91.7 | 183.3 | 210.7 | 366.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 36.7 | 91.7 | 183.3 | 210.7 | 366.7 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 36.7 | 91.7 | 183.3 | 210.7 | 366.7 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
1E4-S | Low Pressure (LP) (with vacuum) | 73.3 | 183.3 | 366.7 | 421.4 | 733.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Low Pressure (LP) (no vacuum) | 73.3 | 183.3 | 366.7 | 421.4 | 733.3 | n/a[4] | n/a[4] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (with vacuum) | 73.3 | 183.3 | 366.7 | 421.4 | 733.3 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| Standard Pressure (SP) (no vacuum) | 73.3 | 183.3 | 366.7 | 421.4 | 733.3 | n/a[8] | n/a[8] | n/a[4] | n/a[4] | n/a[4]
| High Pressure (HP) | 73.3 | 183.3 | 366.7 | 421.4 | 733.3 | n/a[8] | n/a[8] | n/a[8] | n/a[8] | n/a[8]
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range. [8] Operation in this gas is limited to lower pressures due to the gas vapor pressure.
Oxygen Ratio = 0.935 | O2 | Pressure Dependent Flow Rates: Mass Flow Rate (slm @ 0C) WHEN molbloc-S Upstream Pressure is:[1][2][7]
molbloc-S | Pressure Calibration | 20 kPa (3 psia)[5] | 50 kPa (7 psia)[5] | 100 kPa (15 psia)[5] | Minimum without vacuum[3][4] | 200 kPa (30 psia)[4] | 400 kPa (58 psia)[4] | 600 kPa (87 psia)[4] | 700 kPa (100 psia)[4][6] | 1.20 MPa (174 psia)[4][6] | 2 MPa (290 psia)[4][6]
1E1-S | Low Pressure (LP) (with vacuum) | 0.2 | 0.5 | 0.9 | 1.9 | 1.9 | 3.75 | 5.6 | 6.5 | 11.2 | 18.7
| Low Pressure (LP) (no vacuum) | 0.2 | 0.5 | 0.9 | 1.9 | 1.9 | 3.75 | 5.6 | 6.5 | 11.2 | 18.7
| Standard Pressure (SP) (with vacuum) | 0.2 | 0.5 | 0.9 | 1.9 | 1.9 | 3.75 | 5.6 | 6.5 | 11.2 | 18.7
| Standard Pressure (SP) (no vacuum) | 0.2 | 0.5 | 0.9 | 1.9 | 1.9 | 3.75 | 5.6 | 6.5 | 11.2 | 18.7
| High Pressure (HP) | 0.2 | 0.5 | 0.9 | 1.9 | 1.9 | 3.75 | 5.6 | 6.5 | 11.2 | 18.7
2E1-S | Low Pressure (LP) (with vacuum) | 0.4 | 0.9 | 1.9 | 3.4 | 3.7 | 7.475 | 11.25 | 13.1 | 22.5 | 37.4
| Low Pressure (LP) (no vacuum) | 0.4 | 0.9 | 1.9 | 3.4 | 3.7 | 7.475 | 11.25 | 13.1 | 22.5 | 37.4
| Standard Pressure (SP) (with vacuum) | 0.4 | 0.9 | 1.9 | 3.4 | 3.7 | 7.475 | 11.25 | 13.1 | 22.5 | 37.4
| Standard Pressure (SP) (no vacuum) | 0.4 | 0.9 | 1.9 | 3.4 | 3.7 | 7.475 | 11.25 | 13.1 | 22.5 | 37.4
| High Pressure (HP) | 0.4 | 0.9 | 1.9 | 3.4 | 3.7 | 7.475 | 11.25 | 13.1 | 22.5 | 37.4
5E1-S | Low Pressure (LP) (with vacuum) | 0.9 | 2.3 | 4.7 | 7.3 | 9.4 | 18.725 | 28.05 | 32.7 | 56.1 | 93.5
| Low Pressure (LP) (no vacuum) | 0.9 | 2.3 | 4.7 | 7.3 | 9.4 | 18.725 | 28.05 | 32.7 | 56.1 | 93.5
| Standard Pressure (SP) (with vacuum) | 0.9 | 2.3 | 4.7 | 7.3 | 9.4 | 18.725 | 28.05 | 32.7 | 56.1 | 93.5
| Standard Pressure (SP) (no vacuum) | 0.9 | 2.3 | 4.7 | 7.3 | 9.4 | 18.725 | 28.05 | 32.7 | 56.1 | 93.5
| High Pressure (HP) | 0.9 | 2.3 | 4.7 | 7.3 | 9.4 | 18.725 | 28.05 | 32.7 | 56.1 | 93.5
1E2-S | Low Pressure (LP) (with vacuum) | 1.9 | 4.7 | 9.4 | 14.4 | 18.7 | 37.425 | 56.15 | 65.5 | 112.3 | 187.1
| Low Pressure (LP) (no vacuum) | 1.9 | 4.7 | 9.4 | 14.4 | 18.7 | 37.425 | 56.15 | 65.5 | 112.3 | 187.1
| Standard Pressure (SP) (with vacuum) | 1.9 | 4.7 | 9.4 | 14.4 | 18.7 | 37.425 | 56.15 | 65.5 | 112.3 | 187.1
| Standard Pressure (SP) (no vacuum) | 1.9 | 4.7 | 9.4 | 14.4 | 18.7 | 37.425 | 56.15 | 65.5 | 112.3 | 187.1
| High Pressure (HP) | 1.9 | 4.7 | 9.4 | 14.4 | 18.7 | 37.425 | 56.15 | 65.5 | 112.3 | 187.1
2E2-S | Low Pressure (LP) (with vacuum) | 3.7 | 9.4 | 18.7 | 26.4 | 37.4 | 74.825 | 112.25 | 131 | 224.5 | 374.2
| Low Pressure (LP) (no vacuum) | 3.7 | 9.4 | 18.7 | 26.4 | 37.4 | 74.825 | 112.25 | 131 | 224.5 | 374.2
| Standard Pressure (SP) (with vacuum) | 3.7 | 9.4 | 18.7 | 26.4 | 37.4 | 74.825 | 112.25 | 131 | 224.5 | 374.2
| Standard Pressure (SP) (no vacuum) | 3.7 | 9.4 | 18.7 | 26.4 | 37.4 | 74.825 | 112.25 | 131 | 224.5 | 374.2
| High Pressure (HP) | 3.7 | 9.4 | 18.7 | 26.4 | 37.4 | 74.825 | 112.25 | 131 | 224.5 | 374.2
5E2-S | Low Pressure (LP) (with vacuum) | 9.4 | 23.4 | 46.8 | 63.8 | 93.5 | 187.075 | 280.65 | 327.4 | 561.3 | 935.4
| Low Pressure (LP) (no vacuum) | 9.4 | 23.4 | 46.8 | 63.8 | 93.5 | 187.075 | 280.65 | 327.4 | 561.3 | 935.4
| Standard Pressure (SP) (with vacuum) | 9.4 | 23.4 | 46.8 | 63.8 | 93.5 | 187.075 | 280.65 | 327.4 | 561.3 | 935.4
| Standard Pressure (SP) (no vacuum) | 9.4 | 23.4 | 46.8 | 63.8 | 93.5 | 187.075 | 280.65 | 327.4 | 561.3 | 935.4
| High Pressure (HP) | 9.4 | 23.4 | 46.8 | 63.8 | 93.5 | 187.075 | 280.65 | 327.4 | 561.3 | 935.4
1E3-S | Low Pressure (LP) (with vacuum) | 18.7 | 46.8 | 93.5 | 120.4 | 187.1 | 374.175 | 561.25 | 654.8 | 1122.5 | 1870.8
| Low Pressure (LP) (no vacuum) | 18.7 | 46.8 | 93.5 | 120.4 | 187.1 | 374.175 | 561.25 | 654.8 | 1122.5 | 1870.8
| Standard Pressure (SP) (with vacuum) | 18.7 | 46.8 | 93.5 | 120.4 | 187.1 | 374.175 | 561.25 | 654.8 | 1122.5 | 1870.8
| Standard Pressure (SP) (no vacuum) | 18.7 | 46.8 | 93.5 | 120.4 | 187.1 | 374.175 | 561.25 | 654.8 | 1122.5 | 1870.8
| High Pressure (HP) | 18.7 | 46.8 | 93.5 | 120.4 | 187.1 | 374.175 | 561.25 | 654.8 | 1122.5 | 1870.8
2E3-S | Low Pressure (LP) (with vacuum) | 37.4 | 93.5 | 187.1 | 231.9 | 374.2 | 748.35 | 1122.5 | 1309.6 | 2245 | 3741.7
| Low Pressure (LP) (no vacuum) | 37.4 | 93.5 | 187.1 | 231.9 | 374.2 | 748.35 | 1122.5 | 1309.6 | 2245 | 3741.7
| Standard Pressure (SP) (with vacuum) | 37.4 | 93.5 | 187.1 | 231.9 | 374.2 | 748.35 | 1122.5 | 1309.6 | 2245 | 3741.7
| Standard Pressure (SP) (no vacuum) | 37.4 | 93.5 | 187.1 | 231.9 | 374.2 | 748.35 | 1122.5 | 1309.6 | 2245 | 3741.7
| High Pressure (HP) | 37.4 | 93.5 | 187.1 | 231.9 | 374.2 | 748.35 | 1122.5 | 1309.6 | 2245 | 3741.7
5E3-S | Low Pressure (LP) (with vacuum) | 93.5 | 233.9 | 467.7 | 556.6 | 935.4 | 1870.825 | 2806.25 | 3274 | 5612.5 | 9354.2
| Low Pressure (LP) (no vacuum) | 93.5 | 233.9 | 467.7 | 556.6 | 935.4 | 1870.825 | 2806.25 | 3274 | 5612.5 | 9354.2
| Standard Pressure (SP) (with vacuum) | 93.5 | 233.9 | 467.7 | 556.6 | 935.4 | 1870.825 | 2806.25 | 3274 | 5612.5 | 9354.2
| Standard Pressure (SP) (no vacuum) | 93.5 | 233.9 | 467.7 | 556.6 | 935.4 | 1870.825 | 2806.25 | 3274 | 5612.5 | 9354.2
| High Pressure (HP) | 93.5 | 233.9 | 467.7 | 556.6 | 935.4 | 1870.825 | 2806.25 | 3274 | 5612.5 | 9354.2
1E4-S | Low Pressure (LP) (with vacuum) | 187.1 | 467.7 | 935.4 | 1113.1 | 1870.8 | 3741.65 | 5612.5 | 6547.9 | 11225 | 18708.3
| Low Pressure (LP) (no vacuum) | 187.1 | 467.7 | 935.4 | 1113.1 | 1870.8 | 3741.65 | 5612.5 | 6547.9 | 11225 | 18708.3
| Standard Pressure (SP) (with vacuum) | 187.1 | 467.7 | 935.4 | 1113.1 | 1870.8 | 3741.65 | 5612.5 | 6547.9 | 11225 | 18708.3
| Standard Pressure (SP) (no vacuum) | 187.1 | 467.7 | 935.4 | 1113.1 | 1870.8 | 3741.65 | 5612.5 | 6547.9 | 11225 | 18708.3
| High Pressure (HP) | 187.1 | 467.7 | 935.4 | 1113.1 | 1870.8 | 3741.65 | 5612.5 | 6547.9 | 11225 | 18708.3
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.

## Alt_Combination_5 gases
Pressure Dependent Flow Rates: molBloc- S Mass Flow Rate (slm @ 0C):[1][2][7]
Designator / Kf (sccm/kPa)
|  | Pressure Calibration type Min/Max Flow Rates | when Pup is: kPaa | psia | 1E1-S (Kf =10) | 2E1-S (Kf =20) | 5E1-S (Kf =50) | 1E2-S (Kf =100) | 2E2-S (Kf =200) | 5E2-S (Kf =500) | 1E3-S (Kf =1,000) | 2E3-S (Kf =2,000) | 5E3-S (Kf =5,000) | 1E4-S (Kf =10,000)
N2 |  | LP Min[5] | 20 | 3 | 0.2 | 0.4 | 1 | 2 | 4 | 10 | 20 | 40 | 100 | 200
Ratio = | 1 | SP Min[5] | 50 | 7 | 0.5 | 1 | 2.5 | 5 | 10 | 25 | 50 | 100 | 250 | 500
|  | [5] | 100 | 15 | 1 | 2 | 5 | 10 | 20 | 50 | 100 | 200 | 500 | 1000
|  | no vacuum Min[3][4] | (atm on outlet) |  | 2 | 3.5 | 7.7 | 15 | 28 | 67 | 129 | 248 | 596 | 1173
|  | LP Max, HP Min[4] | 200 | 30 | 2 | 4 | 10 | 20 | 40 | 100 | 200 | 400 | 1000 | 2000
|  | SP Max[4] | 600 | 87 | 6 | 12 | 30 | 60 | 120 | 300 | 600 | 1200 | 3000 | 6000
|  | [4][6] | 1200 | 174 | 12 | 24 | 60 | 120 | 240 | 600 | 1200 | 2400 | 6000 | 12000
|  | HP Max[4][6] | 2000 | 290 | 20 | 40 | 100 | 200 | 400 | 1000 | 2000 | 4000 | 10000 | 20000
Air |  | LP Min[5] | 20 | 3 | 0.1966 | 0.3932 | 0.983 | 1.966 | 3.932 | 9.83 | 19.66 | 39.32 | 98.3 | 196.6
Ratio = | 0.983 | SP Min[5] | 50 | 7 | 0.4915 | 0.983 | 2.4575 | 4.915 | 9.83 | 24.575 | 49.15 | 98.3 | 245.75 | 491.5
|  | [5] | 100 | 15 | 0.983 | 1.966 | 4.915 | 9.83 | 19.66 | 49.15 | 98.3 | 196.6 | 491.5 | 983
|  | no vacuum Min[3][4] | (atm on outlet) |  | 2 | 3.4 | 7.6 | 15.2 | 27.4 | 67.1 | 126.7 | 243.9 | 585.2 | 1170.3
|  | LP Max, HP Min[4] | 200 | 30 | 1.966 | 3.932 | 9.83 | 19.66 | 39.32 | 98.3 | 196.6 | 393.2 | 983 | 1966
|  | SP Max[4] | 600 | 87 | 5.898 | 11.796 | 29.49 | 58.98 | 117.96 | 294.9 | 589.8 | 1179.6 | 2949 | 5898
|  | [4][6] | 1200 | 174 | 11.796 | 23.592 | 58.98 | 117.96 | 235.92 | 589.8 | 1179.6 | 2359.2 | 5898 | 11796
|  | HP Max[4][6] | 2000 | 290 | 19.66 | 39.32 | 98.3 | 196.6 | 393.2 | 983 | 1966 | 3932 | 9830 | 19660
Argon |  | LP Min[5] | 20 | 3 | 0.1674 | 0.3348 | 0.837 | 1.674 | 3.348 | 8.37 | 16.74 | 33.48 | 83.7 | 167.4
Ratio = | 0.837 | SP Min[5] | 50 | 7 | 0.4185 | 0.837 | 2.0925 | 4.185 | 8.37 | 20.925 | 41.85 | 83.7 | 209.25 | 418.5
|  | [5] | 100 | 15 | 0.837 | 1.674 | 4.185 | 8.37 | 16.74 | 41.85 | 83.7 | 167.4 | 418.5 | 837
|  | no vacuum Min[3][4] | (atm on outlet) |  | 1.7 | 3 | 6.5 | 12.9 | 23.3 | 57.1 | 107.8 | 207.6 | 498.2 | 996.2
|  | LP Max, HP Min[4] | 200 | 30 | 1.674 | 3.348 | 8.37 | 16.74 | 33.48 | 83.7 | 167.4 | 334.8 | 837 | 1674
|  | SP Max[4] | 600 | 87 | 5.022 | 10.044 | 25.11 | 50.22 | 100.44 | 251.1 | 502.2 | 1004.4 | 2511 | 5022
|  | [4][6] | 1200 | 174 | 10.044 | 20.088 | 50.22 | 100.44 | 200.88 | 502.2 | 1004.4 | 2008.8 | 5022 | 10044
|  | HP Max[4][6] | 2000 | 290 | 16.74 | 33.48 | 83.7 | 167.4 | 334.8 | 837 | 1674 | 3348 | 8370 | 16740
Helium |  | LP Min[5] | 20 | 3 | 0.5294 | 1.0588 | 2.647 | 5.294 | 10.588 | 26.47 | 52.94 | 105.88 | 264.7 | 529.4
Ratio = | 2.647 | SP Min[5] | 50 | 7 | 1.3235 | 2.647 | 6.6175 | 13.235 | 26.47 | 66.175 | 132.35 | 264.7 | 661.75 | 1323.5
|  | [5] | 100 | 15 | 2.647 | 5.294 | 13.235 | 26.47 | 52.94 | 132.35 | 264.7 | 529.4 | 1323.5 | 2647
|  | no vacuum Min[3][4] | (atm on outlet) |  | 9.4 | 13.1 | 25.7 | 51.4 | 91.5 | 199.4 | 398.7 | 695.1 | 1737.8 | 3281
|  | LP Max, HP Min[4] | 200 | 30 | 5.294 | 10.588 | 26.47 | 52.94 | 105.88 | 264.7 | 529.4 | 1058.8 | 2647 | 5294
|  | SP Max[4] | 600 | 87 | 15.882 | 31.764 | 79.41 | 158.82 | 317.64 | 794.1 | 1588.2 | 3176.4 | 7941 | 15882
|  | [4][6] | 1200 | 174 | 31.764 | 63.528 | 158.82 | 317.64 | 635.28 | 1588.2 | 3176.4 | 6352.8 | 15882 | 31764
|  | HP Max[4][6] | 2000 | 290 | 52.94 | 105.88 | 264.7 | 529.4 | 1058.8 | 2647 | 5294 | 10588 | 26470 | 52940
CO2 |  | LP Min[5] | 20 | 3 | 0.159 | 0.318 | 0.795 | 1.59 | 3.18 | 7.95 | 15.9 | 31.8 | 79.5 | 159
Ratio = | 0.795 | SP Min[5] | 50 | 7 | 0.3975 | 0.795 | 1.9875 | 3.975 | 7.95 | 19.875 | 39.75 | 79.5 | 198.75 | 397.5
|  | [5] | 100 | 15 | 0.795 | 1.59 | 3.975 | 7.95 | 15.9 | 39.75 | 79.5 | 159 | 397.5 | 795
|  | no vacuum Min[3][4] | (atm on outlet) |  | 1.4 | 2.5 | 6.2 | 11.1 | 22.1 | 51.2 | 102.4 | 189.3 | 473.2 | 914.1
|  | LP Max, HP Min[4] | 200 | 30 | 1.59 | 3.18 | 7.95 | 15.9 | 31.8 | 79.5 | 159 | 318 | 795 | 1590
|  | SP Max[4] | 600 | 87 | 4.77 | 9.54 | 23.85 | 47.7 | 95.4 | 238.5 | 477 | 954 | 2385 | 4770
|  | [4][6] | 1200 | 174 | 9.54 | 19.08 | 47.7 | 95.4 | 190.8 | 477 | 954 | 1908 | 4770 | 9540
|  | HP Max[4][6] | 2000 | 290 | 15.9 | 31.8 | 79.5 | 159 | 318 | 795 | 1590 | 3180 | 7950 | 15900
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [3] Minimum upstream pressure to achieve critical flow with atmospheric pressure (approximately 100 kPa) downstream of molbloc-S (no vacuum). [4] Assumes proper pressurized gas supply is connected and can be maintained at the given flow rate. [5] Assumes sufficient vacuum source is connected and can be maintained at the given flow rate. [6] Requires Molbox1+S. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.

## Alt_Min Max_with press_5 gases
Pressure Dependent Flow Rates: molbloc- S Mass Flow Rate (slm @ 0C):[1][2][7]
Designator / Kf (sccm/kPa)
[7 empty] | 1E1-S |  | 2E1-S |  | 5E1-S |  | 1E2-S |  | 2E2-S |  | 5E2-S |  | 1E3-S |  | 2E3-S |  | 5E3-S |  | 1E4-S
|  |  | Min Press. |  | Max Press. |  | 1E1-S Kf = | 10 | 1E1-S Kf = | 20 | 1E1-S Kf = | 50 | 1E1-S Kf = | 100 | 1E1-S Kf = | 200 | 1E1-S Kf = | 500 | 1E1-S Kf = | 1000 | 1E1-S Kf = | 2000 | 1E1-S Kf = | 5000 | 1E1-S Kf = | 10000
|  | Pressure Calibration type | kpaa | psia | kPaa | psia | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max
N2 |  | Low Pressure (LP) (with vacuum) | 20 | 3 | 200 | 30 | 0.2 | 2 | 0.4 | 4 | 1 | 10 | 2 | 20 | 4 | 40 | 10 | 100 | 20 | 200 | 40 | 400 | 100 | 1000 | 200 | 2000
Ratio = | 1 | Low Pressure (LP) (no vacuum) | (ATM on Outlet) |  | 200 | 30 | 2 | 2 | 3.5 | 4 | 7.7 | 10 | 15 | 20 | 28 | 40 | 67 | 100 | 129 | 200 | 248 | 400 | 596 | 1000 | 1173 | 2000
|  | Standard Pressure (SP) (with vacuum) | 50 | 7 | 600 | 87 | 0.5 | 6 | 1 | 12 | 2.5 | 30 | 5 | 60 | 10 | 120 | 25 | 300 | 50 | 600 | 100 | 1200 | 250 | 3000 | 500 | 6000
|  | Standard Pressure (SP) (no vacuum) | (ATM on Outlet) |  | 600 | 87 | 2 | 6 | 3.5 | 12 | 7.7 | 30 | 15 | 60 | 28 | 120 | 67 | 300 | 129 | 600 | 248 | 1200 | 596 | 3000 | 1173 | 6000
|  | High Pressure (HP) | 200 | 30 | 2000 | 290 | 2 | 20 | 4 | 40 | 10 | 100 | 20 | 200 | 40 | 400 | 100 | 1000 | 200 | 2000 | 400 | 4000 | 1000 | 10000 | 2000 | 20000
Air |  | Low Pressure (LP) (with vacuum) | 20 | 3 | 200 | 30 | 0.1966 | 1.966 | 0.3932 | 3.932 | 0.983 | 9.83 | 1.966 | 19.66 | 3.932 | 39.32 | 9.83 | 98.3 | 19.66 | 196.6 | 39.32 | 393.2 | 98.3 | 983 | 196.6 | 1966
Ratio = | 0.983 | Low Pressure (LP) (no vacuum) | (ATM on Outlet) |  | 200 | 30 | 2 | 1.966 | 3.4 | 3.932 | 7.6 | 9.83 | 15.2 | 19.66 | 27.4 | 39.32 | 67.1 | 98.3 | 126.7 | 196.6 | 243.9 | 393.2 | 585.2 | 983 | 1170.3 | 1966
|  | Standard Pressure (SP) (with vacuum) | 50 | 7 | 600 | 87 | 0.4915 | 5.898 | 0.983 | 11.796 | 2.4575 | 29.49 | 4.915 | 58.98 | 9.83 | 117.96 | 24.575 | 294.9 | 49.15 | 589.8 | 98.3 | 1179.6 | 245.75 | 2949 | 491.5 | 5898
|  | Standard Pressure (SP) (no vacuum) | (ATM on Outlet) |  | 600 | 87 | 2 | 5.898 | 3.4 | 11.796 | 7.6 | 29.49 | 15.2 | 58.98 | 27.4 | 117.96 | 67.1 | 294.9 | 126.7 | 589.8 | 243.9 | 1179.6 | 585.2 | 2949 | 1170.3 | 5898
|  | High Pressure (HP) | 200 | 30 | 2000 | 290 | 1.966 | 19.66 | 3.932 | 39.32 | 9.83 | 98.3 | 19.66 | 196.6 | 39.32 | 393.2 | 98.3 | 983 | 196.6 | 1966 | 393.2 | 3932 | 983 | 9830 | 1966 | 19660
Argon |  | Low Pressure (LP) (with vacuum) | 20 | 3 | 200 | 30 | 0.1674 | 1.674 | 0.3348 | 3.348 | 0.837 | 8.37 | 1.674 | 16.74 | 3.348 | 33.48 | 8.37 | 83.7 | 16.74 | 167.4 | 33.48 | 334.8 | 83.7 | 837 | 167.4 | 1674
Ratio = | 0.837 | Low Pressure (LP) (no vacuum) | (ATM on Outlet) |  | 200 | 30 | 1.7 | 1.674 | 3 | 3.348 | 6.5 | 8.37 | 12.9 | 16.74 | 23.3 | 33.48 | 57.1 | 83.7 | 107.8 | 167.4 | 207.6 | 334.8 | 498.2 | 837 | 996.2 | 1674
|  | Standard Pressure (SP) (with vacuum) | 50 | 7 | 600 | 87 | 0.4185 | 5.022 | 0.837 | 10.044 | 2.0925 | 25.11 | 4.185 | 50.22 | 8.37 | 100.44 | 20.925 | 251.1 | 41.85 | 502.2 | 83.7 | 1004.4 | 209.25 | 2511 | 418.5 | 5022
|  | Standard Pressure (SP) (no vacuum) | (ATM on Outlet) |  | 600 | 87 | 1.7 | 5.022 | 3 | 10.044 | 6.5 | 25.11 | 12.9 | 50.22 | 23.3 | 100.44 | 57.1 | 251.1 | 107.8 | 502.2 | 207.6 | 1004.4 | 498.2 | 2511 | 996.2 | 5022
|  | High Pressure (HP) | 200 | 30 | 2000 | 290 | 1.674 | 16.74 | 3.348 | 33.48 | 8.37 | 83.7 | 16.74 | 167.4 | 33.48 | 334.8 | 83.7 | 837 | 167.4 | 1674 | 334.8 | 3348 | 837 | 8370 | 1674 | 16740
Helium |  | Low Pressure (LP) (with vacuum) | 20 | 3 | 200 | 30 | 0.5294 | 5.294 | 1.0588 | 10.588 | 2.647 | 26.47 | 5.294 | 52.94 | 10.588 | 105.88 | 26.47 | 264.7 | 52.94 | 529.4 | 105.88 | 1058.8 | 264.7 | 2647 | 529.4 | 5294
Ratio = | 2.647 | Low Pressure (LP) (no vacuum) | (ATM on Outlet) |  | 200 | 30 | NA | NA | NA | NA | 25.7 | 26.47 | 51.4 | 52.94 | 91.5 | 105.88 | 199.4 | 264.7 | 398.7 | 529.4 | 695.1 | 1058.8 | 1737.8 | 2647 | 3281 | 5294
|  | Standard Pressure (SP) (with vacuum) | 50 | 7 | 600 | 87 | 1.3235 | 15.882 | 2.647 | 31.764 | 6.6175 | 79.41 | 13.235 | 158.82 | 26.47 | 317.64 | 66.175 | 794.1 | 132.35 | 1588.2 | 264.7 | 3176.4 | 661.75 | 7941 | 1323.5 | 15882
|  | Standard Pressure (SP) (no vacuum) | (ATM on Outlet) |  | 600 | 87 | NA | NA | NA | NA | 25.7 | 79.41 | 51.4 | 158.82 | 91.5 | 317.64 | 199.4 | 794.1 | 398.7 | 1588.2 | 695.1 | 3176.4 | 1737.8 | 7941 | 3281 | 15882
|  | High Pressure (HP) | 200 | 30 | 2000 | 290 | 5.294 | 52.94 | 10.588 | 105.88 | 26.47 | 264.7 | 52.94 | 529.4 | 105.88 | 1058.8 | 264.7 | 2647 | 529.4 | 5294 | 1058.8 | 10588 | 2647 | 26470 | 5294 | 52940
CO2 |  | Low Pressure (LP) (with vacuum) | 20 | 3 | 200 | 30 | 0.159 | 1.59 | 0.318 | 3.18 | 0.795 | 7.95 | 1.59 | 15.9 | 3.18 | 31.8 | 7.95 | 79.5 | 15.9 | 159 | 31.8 | 318 | 79.5 | 795 | 159 | 1590
Ratio = | 0.795 | Low Pressure (LP) (no vacuum) | (ATM on Outlet) |  | 200 | 30 | 1.4 | 1.59 | 2.5 | 3.18 | 6.2 | 7.95 | 11.1 | 15.9 | 22.1 | 31.8 | 51.2 | 79.5 | 102.4 | 159 | 189.3 | 318 | 473.2 | 795 | 914.1 | 1590
|  | Standard Pressure (SP) (with vacuum) | 50 | 7 | 600 | 87 | 0.3975 | 4.77 | 0.795 | 9.54 | 1.9875 | 23.85 | 3.975 | 47.7 | 7.95 | 95.4 | 19.875 | 238.5 | 39.75 | 477 | 79.5 | 954 | 198.75 | 2385 | 397.5 | 4770
|  | Standard Pressure (SP) (no vacuum) | (ATM on Outlet) |  | 600 | 87 | 1.4 | 4.77 | 2.5 | 9.54 | 6.2 | 23.85 | 11.1 | 47.7 | 22.1 | 95.4 | 51.2 | 238.5 | 102.4 | 477 | 189.3 | 954 | 473.2 | 2385 | 914.1 | 4770
|  | High Pressure (HP) | 200 | 30 | 2000 | 290 | 1.59 | 15.9 | 3.18 | 31.8 | 7.95 | 79.5 | 15.9 | 159 | 31.8 | 318 | 79.5 | 795 | 159 | 1590 | 318 | 3180 | 795 | 7950 | 1590 | 15900
[1] Flow values in table are valid only when critical flow is established. [2] When volumetrically based flow units with reference temperatures other than 0 ˚C are used, flow values will generally be higher: the flow values for a given molbloc upstream pressure are approximately 7.7% higher when expressed in slm @ 21.1 ˚C (same as 70 ˚F). Flow values at a given pressure may vary by up to ±2% due to flow path machining tolerances. [7] DUTs or other restrictions downstream of the molbloc-S can also limit useable flow range.