## Leak Test
700HPPK Leak Test Worksheet (includes 700HPP & 700HPM)
Cells with boxes around them can be changed or have values entered in them. You can copy the results in cells A28 and A29 and paste them into a service report.
A. You will need a digital pressure gauge that measures up to at least 3300 psi (23 MPa or 230 bar) and has at least 0.1 psi (0.001 MPa or 0.01 bar) resolution
i. Generate and apply 3300 psi (23 MPa or 230 bar) to the component(s) to be tested. Or apply the maximum pressure of the lowest rated component.
ii. If not testing the pressure generation source, isolate it from the component(s) being tested.
iii. Wait (dwell) for 4 minutes for any temperature and other effects to stabilize
Leak rate limit = | 20 | ppm/sec
| 0.002 | % of reading
Units
Start pressure (Ps) = | 3001.65 | psi | psi
bar
Final pressure (Pf) = | 2999.76 | psi | MPa
kPa
Time of leak test = | 60 | sec | kpsi
| 1 | min
Average pressure (Pavg) = | 3000.705 | psi
(Pf - Ps) / 2
Leak rate limit = | 3.6008 | psi/min | Conditional Text
is less than the Leak rate limit of 20 ppm/sec.
Leak rate = (Pf – Ps) / 1 min = | -1.89 | psi/min | is less than the Leak rate limit of 20 ppm/sec.
= (Leak rate / Pavg) *1E6 | -10.4975 | ppm/sec | is less than the Leak rate limit of 20 ppm/sec.
is less than the Leak rate limit of 20 ppm/sec.
The Leak rate of -10.5 ppm/sec at 3000.705 psi is less than the Leak rate limit of 20 ppm/sec. |  |  | Was the test run correctly? Something isn't working.
The Leak test result is PASS. |  |  | PASS.
PASS.

## VV Test
700HPM Performance Test Worksheet
You will need a digital pressure gauge that measures up to at least 3000 psi (20.7 MPa or 206.9 bar) and has at least 0.1 psi (0.001 MPa or 0.01 bar) resolution.
You will also need a pressure generation source that can supply up to at least 3000 psi (20.7 MPa or 206.9 bar).
This would typically be the 700HPP (high pressure pump).
Cells with boxes around them can be changed or have values entered in them.
A typical note for a service report would be "700HPM variable volume checked through its range for smooth operation and correct pressure changes."
1 | Open the INLET valve on the 700HPM and turn the variable volume knob (VV) fully out, all the way counter-clockwise (CCW).
2. | Generate and apply 85 psi (0.586 MPa or 0.59 bar) to the 700HPM. Do not use the VV to adjust the pressure. Leave it fully in.
3. | Close the INLET valve. The pressure will raise a little bit.
4. | Wait for the pressure to stabilize (stop changing) and record the pressure.
| Start pressure = | 85.1 | psi
5. | Turn the VV fully in (CW), wait for the pressure to stabilize and record the pressure.
| Final pressure = | 107.5 | psi
| Pressure difference = | 22.4 | psi | % change = | 26.32%
The pressure difference will vary based on what is connected to the gauge and DUT (Device Under Test) ports of the 700HPM.
With two low volume gauges, the pressure difference should be about 25%, or 21.28 psi
With one low volume gauge and a higher volume gauge, the pressure difference should be about 15% to 20%, or 12.77 to 17.02 psi
6. | In step 5, the movement of the VV should be smooth without any irregularities. If not, the plunger seal should be replaced with a SK-700HPM-FS seal kit.
7. | Open the INLET valve and with the VV still fully in (CW), generate and apply 1100 psi (7.58 MPa or 75.8 bar) to the 700HPM.
Do not use the VV to adjust the pressure. Leave it fully in.
8. | Close the INLET valve. The pressure will raise a little bit.
9. | Wait for the pressure to stabilize (stop changing) and record the pressure.
| Start pressure = | 1124.8 | psi
10. | Turn the VV fully out (CCW), wait for the pressure to stabilize and record the pressure.
| Final pressure = | 854.1 | psi
| Pressure difference = | -270.7 | psi | % change = | -24.07%
The pressure difference will vary based on what is connected to the gauge and DUT (Device Under Test) ports of the 700HPM.
With two low volume gauges, the pressure difference should be about -25%, or -281.20 psi
With one low volume gauge and a higher volume gauge, the pressure difference should be about -15% to -20%, or -168.72 to -224.96 psi
11. | Turn the VV in (CW) until it is at mid-stroke (1 inch or 2.5 cm). The pressure should be around 1000 psi.
12. | Wait for the pressure to stabilize (stop changing) and record the pressure.
| Start pressure = | 998.3 | psi
13. | Turn the VV one turn in (CW), wait for the pressure to stabilize and record the pressure.
| Final pressure = | 1004.4 | psi
| Pressure difference = | 6.1 | psi | % change = | 0.611%
The pressure difference will vary based on what is connected to the gauge and DUT (Device Under Test) ports of the 700HPM.
With two low volume gauges, the pressure difference should be about 0.5%, or 5.02 psi
With one low volume gauge and a higher volume gauge, the pressure difference should be about 0.4%, or 3.99 psi
14. | In steps 10 and 13, the movement of the VV should be smooth without any irregularities. If not, the plunger seal should be replaced with a SK-700HPM-FS seal kit."
15. | Open the INLET valve and turn the VV fully out (CCW).
With the VV fully out (CCW), generate and apply 3000 psi (20.7 MPa or 206.9 bar) to the 700HPM.
Do not use the VV to adjust the pressure. Leave it fully out.
16. | Close the INLET valve. The pressure will raise a little bit.
17. | Wait for the pressure to stabilize (stop changing) and record the pressure.
| Start pressure = | 3095.9 | psi
18. | Turn the VV fully out (CCW), wait for the pressure to stabilize and record the pressure.
| Final pressure = | 2351.8 | psi
| Pressure difference = | -744.1 | psi | % change = | -24.04%
The pressure difference will vary based on what is connected to the gauge and DUT (Device Under Test) ports of the 700HPM.
With two low volume gauges, the pressure difference should be about -25%, or -773.98 psi
With one low volume gauge and a higher volume gauge, the pressure difference should be about -15% to -20%, or -464.39 to -619.18 psi
19. | Turn the VV in (CW) until it is at mid-stroke (1 inch or 2.5 cm). The pressure should be around 2725 psi.
20. | Wait for the pressure to stabilize (stop changing) and record the pressure.
| Start pressure = | 2724.7 | psi
21. | Turn the VV one turn out (CCW), wait for the pressure to stabilize and record the pressure.
| Final pressure = | 2715.1 | psi
| Pressure difference = | -9.6 | psi | % change = | -0.3523%
The pressure difference will vary based on what is connected to the gauge and DUT (Device Under Test) ports of the 700HPM.
With two low volume gauges, the pressure difference should be about -0.5%, or -13.58 psi
With one low volume gauge and a higher volume gauge, the pressure difference should be about -0.4%, or -10.90 psi
22. | In steps 18 and 21, the movement of the VV should be smooth without any irregularities. If not, the plunger seal should be replaced with a SK-700HPM-FS seal kit."