## Coefficient Calcs
Slope / Offset calculation tutorial
As Found |  |  |  | Backed Out Coefs: |  | Predicted As Left Data |  | Actual As Left Data |  |  | Aggreement
Point | Reference Pressure | DUT Pressure | Error (DUT-Ref) | Reference Pressure | Raw DUT Pressure | Corrected DUT Pressure | Predicted A/L Error | Reference Pressure | DUT Pressure | A/L Error | between Predicted vs Actual A/L | Tolerance(+) | Tolerance(-)
| psi | psi | psi | psi | psi | psi | psi | psi | psi | psi | psi | psi | psi
1.1 | 20 | 19.973 | -0.027 | 20 | 19.6581 | 20.0108 | 0.010839 | 20 | 19.99 | -0.01 | 0.020839 | 0.015 | -0.015
1.2 | 100 | 99.973 | -0.027 | 100 | 99.6786 | 100.0115 | 0.011452 | 100 | 99.986 | -0.014 | 0.025452 | 0.015 | -0.015
1.3 | 200 | 199.967 | -0.033 | 200 | 199.6982 | 200.0062 | 0.0062178 | 200 | 200.001 | 0.001 | 0.0052178 | 0.02 | -0.02
1.4 | 300 | 299.958 | -0.042 | 300 | 299.7148 | 299.998 | -0.0020162 | 300 | 300.002 | 0.002 | -0.0040162 | 0.03 | -0.03
1.5 | 400 | 399.95 | -0.05 | 400 | 399.7325 | 399.9907 | -0.0092502 | 400 | 400.005 | 0.005 | -0.01425 | 0.04 | -0.04
1.6 | 500 | 499.974 | -0.026 | 500 | 499.7821 | 500.0155 | 0.015516 | 500 | 500.006 | 0.006 | 0.009516 | 0.05 | -0.05
1.7 | 400 | 399.95 | -0.05 | 400 | 399.7325 | 399.9907 | -0.0092502 | 400 | 399.987 | -0.013 | 0.0037498 | 0.04 | -0.04
1.8 | 300 | 299.956 | -0.044 | 300 | 299.7128 | 299.996 | -0.0040162 | 300 | 299.985 | -0.015 | 0.010984 | 0.03 | -0.03
1.9 | 200 | 199.959 | -0.041 | 200 | 199.6902 | 199.9982 | -0.0017822 | 200 | 199.996 | -0.004 | 0.0022178 | 0.02 | -0.02
1.1 | 100 | 99.956 | -0.044 | 100 | 99.6616 | 99.9945 | -0.0055483 | 100 | 99.995 | -0.005 | -0.0005483 | 0.015 | -0.015
1.11 | 20 | 19.95 | -0.05 | 20 | 19.6351 | 19.9878 | -0.012161 | 20 | 20.001 | 0.001 | -0.013161 | 0.015 | -0.015
Calibration Coefficients: if the previous coefficients are not known, then use 0 and 1:
0.31994 | As Found Intercept coefficient (PA or C0 coefficient)
0.99974 | As Found Slope coefficient (PM or C1 coefficient)
0 | As Found AutoZero value, if applicable (AutoZero or zOffset coefficient)
0.35763 | As Left Intercept coefficient (PA or C0 coefficient)
0.99975 | As Left Slope coefficient (PM or C1 coefficient)
0 | As Left AutoZero value (should always be zero for absolute sensors)
s
The "Backed Out Coefs" table shows what the DUT data would look like with the As Found coefficients removed (raw data from the device). The linear regression is performed on this raw data. Based on the new slope and intercept values, a prediction of what the DUT data could (would) look like is given in the "Predicted As Left data" table.
The coefficents in the green cells are calculated and should be entered into the device before performing the Verification Of Calibration (VOC) run (As Left data).
Cell B27 has a field for an “AutoZero” value. This is for instruments that use a third coefficient for AutoZero adjustments. This would be known as “zOffset”, or “AutoZ” and it is commonly used with Fluke Calibration products. If an AutoZero value is not known or not used, then use a value of “0” in this cell.
The yellow shaded fields are user-input. This is where the As Found calibration data must be entered in, and the As Found coefficients must be entered in the yellow cells below. Use 0 for intercept and 1 for slope if these are unknown.
Note on K8: Daniels, Matt: type in test data from an As Left verification to see how it compares to the predicted As Left results.
Note on L8: Daniels, Matt: type in test data from an As Left verification to see how it compares to the predicted As Left results.