## Unccalc1
Limits | [8 empty] | System Uncertainty
Max Pressure Value | 350000 | Pa | 50.7632 | psi | PABS | 350000 | Enter Min P (Pa) | 10000 |  | Pressure (Pa) | U(95) (Pa) | U(95) (Pa)
Mode | 1 | 1 for ABS, 2 for Gauge, 3 for ABS by ATM |  |  |  |  | psi | 1.4504 | Max | 350000 | 3.902 | 11.1485 | Absolute by application of vacuum mode
Media :gas | 1 | 1 for gas, 2 for oil, 3 for water (must be 1 if mode is 1) |  |  |  |  | key |  | Mid | 175000 | 1.9826 | 11.3294
KN (mass to pressure) | 10001.5196 | Pa/kg | 1.4506 | psi/kg |  | Entry Field |  |  | Min | 10000 | 0.43577 | 43.577
Oil Density | 916 | kg/m3 |  |  |  | Not Used
Maximum Head Height | 0.1 | m | 0.3937 | inches |  | Password | Fluke
Reference Temperature | 23 | °C What temperature is the effective area given | [5 empty] | psi to Pa conv. |  | 0.0010% of rdg | + 0.29 Pa
P-C Temperature Max Limit | 3 | °C What is the maximum difference in p-c temperature and reference temperature | [5 empty] | 50
Confidence | 95 | % | [5 empty] | 344737.9543
Influence Quantities |  |  |  | Influence Uncertainties | [5 empty] | Pressure Variant Terms
Influence | Typical | Approx. |  | Uncertainty Value |  | Type Value | Distribution | Uncertainty | Equiv. Std. | Sensitivity | Sensitivity | Std. Unc. | DOF | Variance | Expanded
Description | Symbol | Value | Units | Either % or value in unit 2σ or 95% |  | 1=relative(%), 2=constant | 1=normal, 2=rectangular |  | Unc. (1s) | Coefficient | Value | (Pa) | ----- | (Pa)² | (2σ, PPM)
Effective Area Terms
Effective Area | AE(Tref,0) | 0.00098052 | m2 | 0.0009 | % | 1 | 1 | 8.8246e-09 | 4.4123e-09 | Ao-1 x P | 356954909.4558 | 1.575 | 30 | 2.4806 | 9
Piston Thermal Coef. | aP | 4.55e-06 | °C-1 | 5 | % | 1 | 2 | 2.275e-07 | 1.3135e-07 | (TMAX - TREF) x P | 1050000 | 0.13791 | 30 | 0.01902 | 0.78808
Cylinder Thermal Coef. | aC | 4.55e-06 | °C-1 | 5 | % | 1 | 2 | 2.275e-07 | 1.3135e-07 | (TMAX - TREF) x P | 1050000 | 0.13791 | 30 | 0.01902 | 0.78808
Piston-Cylinder Temp. | T | 26 | °C | 0.1 | °C | 2 | 1 | 0.1 | 0.05 | (αp+αc) x P | 3.185 | 0.15925 | 30 | 0.025361 | 0.91
1st Order Pressure Coef. | l,b1 | 4.2e-06 | MPa-1 | 0 | % | 1 | 1 | 0 | 0 | P2 | 122500000000 | 0 | 30 | 0 | 0
2nd Order Pressure Coef. | b2 | 0 | MPa-2 | 0 | % | 1 | 1 | 0 | 0 | P3 | 42875000000000000 | 0 | 30 | 0 | 0
Linearity | P | 350000 | Pa | 0.0002 | % | 1 | 1 | 0.7 | 0.35 | 1 | 1 | 0.35 | 30 | 0.1225 | 2
Sensitivity/Performance | P | 350000 | Pa | 1.5e-05 | % | 1 | 1 | 0.0525 | 0.02625 | 1 | 1 | 0.02625 | 30 | 0.00068906 | 0.15
Sensitivity Limit | P | 350000 | Pa | 0.015 | Pa | 2 | 1 | 0.015 | 0.0075 | 1 | 1 | 0.0075 | 30 | 5.625e-05 | 0.042857
Ao Stability | Ao | 0.00098052 | m2 | 0 | % | 1 | 2 | 0 | 0 | Ao-1 x P | 356954909.4558 | 0 | 30 | 0 | 0
Force Terms
Mass Load | M | 34.9947 | kg | 0.0005 | % | 1 | 1 | 0.00017497 | 8.7487e-05 | gl/Ae | 10001.5196 | 0.875 | 30 | 0.76562 | 5
Mass Stability | M | 34.9947 | kg | 0 | % | 1 | 2 | 0 | 0 | gl/Ae | 10001.5196 | 0 | 30 | 0 | 0
AMH Lubrication | M | 1.5e-05 | kg | 1e-05 | kg | 2 | 1 | 1e-05 | 5e-06 | gl/Ae | 10001.5196 | 0.050008 | 30 | 0.0025008 | 0.28576
Piston Mass | M | 0.2 | kg | 3e-06 | kg | 2 | 1 | 3e-06 | 1.5e-06 | gl/Ae | 10001.5196 | 0.015002 | 30 | 0.00022507 | 0.085727
Bell Mass | M | 0.3 | kg | 5e-06 | kg | 2 | 1 | 5e-06 | 2.5e-06 | gl/Ae | 10001.5196 | 0.025004 | 30 | 0.00062519 | 0.14288
Local Gravity | gL | 9.8066 | m/sec2 | 0.0002 | % | 1 | 1 | 1.9613e-05 | 9.8066e-06 | gL-1 x P | 35690.0675 | 0.35 | 30 | 0.1225 | 2
Air Density | rAIR | 1.2 | kg/m3 | 0 | kg/m3 | 2 | 1 | 0 | 0 | rM-1 x P | 50 | 0 | 30 | 0 | 0
Ambient Pressure | PAMB | 97000 | Pa | 140 | Pa | 2 | 1 | 140 | 0 | 1.5E-9 x P | 0.525 | 0 | 30 | 0 | 0
Ambient Temperature | TAMB | 23 | °C | 1 | °C | 2 | 1 | 1 | 0 | 5.5E-7 x P | 0.1925 | 0 | 30 | 0 | 0
Ambient Humidity | HAMB | 50 | %RH | 10 | %RH | 2 | 1 | 10 | 0 | 1.2E-8 x P | 0.0042 | 0 | 30 | 0 | 0
Mass Load AVG Density | rM | 7000 | kg/m3 | 40 | kg/m3 | 2 | 2 | 40 | 23.094 | rAIR/rM2 x P | 0.0085714 | 0.19795 | 30 | 0.039184 | 1.1311
Surface Tension | τ | 0.032 | N/m | 10 | % | 1 | 2 | 0.0032 | 0 | πD/A0 | 113.2081 | 0 | 30 | 0 | 0
Verticality (<0.5 deg.) | q | 0 | deg. | 0.167 | deg. | 2 | 1 | 0.167 | 0.0835 | 1 - COSΦ x P | 0.37168 | 0.37168 | 30 | 0.13814 | 2.1239
System Terms
Density of Fluid | rFLUID | 4.0231 | kg/m3 | 0.1 | % | 1 | 1 | 0.0040231 | 0.0020116 | Aoh/M x P | 0.98067 | 0.0019727 | 30 | 3.8915e-06 | 0.011272
Fluid Head Height | h | 0.1 | m | 0.01 | m | 2 | 1 | 0.01 | 0.005 | g(rFLUID-rAIR) | 39.4536 | 0.19727 | 30 | 0.038915 | 1.1272
Float Position | fpi | 0.007 | m | 0.0001 | m | 2 | 1 | 0.0001 | 5e-05 | g(rFLUID-rAIR) | 39.4536 | 0.0019727 | 30 | 3.8915e-06 | 0.011272
Barometer | Pbaro | 97000 | Pa | 10 | Pa | 2 | 1 | 10 | 0 | 1 | 1 | 0 | 30 | 0 | 0
Reference Pressure | PREF | 4 | Pa | 10 | % | 1 | 1 | 0.4 | 0.2 | 1 | 1 | 0.2 | 30 | 0.04 | 1.1429
Additional | P | 350000 | Pa | 0 | Pa | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 30 | 0 | 0
This Spreadsheet is provided at no cost and was developed primarily as a learning tool. Fluke Corporation does not guarantee the correctness of the calculations here and does not claim any responsibility for its use. For questions or comments on this spreadsheet please contact Mike Bair at mike.bair@flukecal.com. | [10 empty] | Combined Variance: | 3.815 | Pa |  | Coverage
[11 empty] | Combined Uncertainty (k=1): | 1.9532 | Pa | k | 1.9977
[11 empty] | Expanded Uncertainty (95%): | 3.902 | Pa | DOF | 64.2589
[11 empty] | Expanded Uncertainty (95%): | 11.1485 | ppm
Note on B2: Enter maximum pressure in Pa.
Note on I2: Enter the minimum pressure in Pa that you will measure with the Piston Gauge
Note on B7: Enter the maximum head height made with the Piston Gauge.
Note on B8: Enter the p-c effective area reference temperature from the calibration certificate.
Note on B9: Enter the largest temperature correction made for temperature of the piston-cylinder. It is fine to be conservative here. If your lab is 20 to 26 deg C and reference temperature is 23, then this value is 3.
Note on I9: You can use this psi to Pa converter if you are used to using psi. For other units you can just use a simple unit converter tool.
Note on B10: Enter the confidence for the complete uncertainty analysis.
Note on N12: Note that entering 30 for these gives a k of 2. If you are less confident of these uncertainties then reduce these values as appropriate.
Note on E13: Enter the value of the uncertainty for each parameter. It can be a % of the Approx Value column or a value in the unit of measure for that parameter. Enter 1 in the column to the right for % of the parameter value or 2 if it is in the parameter's unit.
Note on C15: Enter the effective area value from the calibration certificate in M^2.
Note on C16: Enter the thermal expansion coefficient for the piston, then cylinder below. If you only know the combination of them you can enter here and use 0 for cylinder.
Note on C18: The reference value plus maximum difference from above is used here. So it is locked.
Note on C19: Enter the pressure distortion coefficient from the certificate here and if there is a 2nd order in the cell below. This must be in 1/MPa. If the report is in another unit use the inverse of the unit conversion. For the 2nd order coefficient use the inverse of the unit conversion squared.
Note on E19: You should only include this uncertainty if you are using a theoretical deformation coefficient as with many PG7000 Piston Gauges.
Note on C21: Always the pressure here so is just what is entered above.
Note on E21: Enter from the PG uncertainty technical Note
Note on E22: You can enter sensitivity in % of value here. Just below you can enter the threshold value. You can find these from the uncertainty technical notes
Note on E24: If your effective area uncertaitny does not inlcude drift you can enter it here separately
Note on E27: Enter a value for stability here if not included in mass uncertainty above.
Note on E28: Enter uncertainty due to AMH lubrication grease on threads of stem and binary tray from PG9000 technical note.
Note on E29: If piston (or cylinder if floating piece) mass uncertainty is different from main mass enter here.
Note on E30: If bell uncertainty is different from main mass then enter here.
Note on C32: You can enter the air density here. Nominal is ok. The spreadsheet is set up so that you can just enter an overall uncertainty for air density, or you can enter the uncertainty for the ambient conditions used to calculate air density (second may be easier). If you enter zero for Air Density, ambient data is used. If you enter an air density uncertainty then ambient data is not use as this would be doubling the uncertainty.
Note on E32: If this is grayed out and not used you have selected absolte by vac mode.
Note on C36: You should enter worse case average density here. The lowest value is the biggest correction and the highest uncertainty.
Note on E36: It is not intuitive but the uncertainty in mass density gets included when in absolute by vac mode, not gauge.
Note on E38: Convert verticality in minutes to degrees to enter here (minutes/60)
Note on C44: This is limit. Enter the highest vacuum pressure that will be used. If you always wait until vacuum is less than 2 Pa, then enter 2.
Note on C45: This is a wild card uncertianty if there is one you want to inlcude that is not in the list here.

## Unccalc2 (hidden)
Limits
Pressure Value | 175000 | Pa | 25.3816 | psi | PABS | 175000
Mode | 1 | 1 for ABS, 2 for Gauge, 3 for ABS by ATM
Media :gas | 1 | 1 for gas, 2 for oil, 3 for water |  |  |  | Entry Field
KN (mass to pressure) | 10001.5196 | Pa/kg | 1.4506 | psi/kg |  | Not Used
Oil Density | 916 | kg/m3 |  |  |  | Password | Fluke
Maximum Head Height | 0.1 | m | 0.3937 | inches
Reference Temperature | 23 | °C What temperature is the effective area given
P-C Temperature Max Limit | 3 | °C What is the maximum difference in room temperature and reference temperature
Confidence | 95 | %
Influence Quantities |  |  |  | Influence Uncertainties | [5 empty] | Pressure Variant Terms
Influence | Typical | Approx. |  | Uncertainty Value |  | Type Value | Distribution | Uncertainty | Equiv. Std. | Sensitivity | Sensitivity | Std. Unc. | DOF | Variance | Expanded
Description | Symbol | Value | Units | Either % or value in unit 2σ or 95% |  | 1=relative(%), 2=constant | 1=normal, 2=rectangular |  | Unc. (1s) | Coefficient | Value | (Pa) | ----- | (Pa)² | (2σ, PPM)
Effective Area Terms
Effective Area | Ao | 0.00098052 | m2 | 0.0009 | % | 1 | 1 | 8.8246e-09 | 4.4123e-09 | Ao-1 x P | 178477454.7279 | 0.7875 | 30 | 0.62016 | 9
Piston Thermal Coef. | aP | 4.55e-06 | °C-1 | 5 | % | 1 | 2 | 2.275e-07 | 1.3135e-07 | (TMAX - TREF) x P | 525000 | 0.068957 | 30 | 0.0047551 | 0.78808
Cylinder Thermal Coef. | aC | 4.55e-06 | °C-1 | 5 | % | 1 | 2 | 2.275e-07 | 1.3135e-07 | (TMAX - TREF) x P | 525000 | 0.068957 | 30 | 0.0047551 | 0.78808
Piston-Cylinder Temp. | T | 26 | °C | 0.1 | °C | 2 | 1 | 0.1 | 0.05 | (αp+αc) x P | 1.5925 | 0.079625 | 30 | 0.0063401 | 0.91
1st Order Pressure Coef. | l,b1 | 4.2e-06 | MPa-1 | 0 | % | 1 | 1 | 0 | 0 | P2 | 30625000000 | 0 | 30 | 0 | 0
2nd Order Pressure Coef. | b2 | 0 | MPa-2 | 0 | % | 1 | 1 | 0 | 0 | P3 | 5359375000000000 | 0 | 30 | 0 | 0
Linearity | P | 175000 | Pa | 0.0002 | % | 1 | 1 | 0.35 | 0.175 | 1 | 1 | 0.175 | 30 | 0.030625 | 2
Performance | P | 175000 | Pa | 1.5e-05 | % | 1 | 1 | 0.02625 | 0.013125 | 1 | 1 | 0.013125 | 30 | 0.00017227 | 0.15
Sensitivity Limit | P | 175000 | Pa | 0.015 | Pa | 2 | 1 | 0.015 | 0.0075 | 1 | 1 | 0.0075 | 30 | 5.625e-05 | 0.085714
Ao Stability (2 yrs) | Ao | 0.00098052 | m2 | 0 | % | 1 | 2 | 0 | 0 | Ao-1 x P | 178477454.7279 | 0 | 30 | 0 | 0
Force Terms
Mass Load | M | 17.4973 | kg | 0.0005 | % | 1 | 1 | 8.7487e-05 | 4.3743e-05 | gl/Ae | 10001.5196 | 0.4375 | 30 | 0.19141 | 5
Mass Stability | M | 17.4973 | kg | 0 | % | 1 | 2 | 0 | 0 | gl/Ae | 10001.5196 | 0 | 30 | 0 | 0
AMH Lubrication | M | 1.5e-05 | kg | 1e-05 | kg | 2 | 1 | 1e-05 | 5e-06 | gl/Ae | 10001.5196 | 0.050008 | 30 | 0.0025008 | 0.57152
Piston Mass | M | 0.2 | kg | 3e-06 | kg | 2 | 1 | 3e-06 | 1.5e-06 | gl/Ae | 10001.5196 | 0.015002 | 30 | 0.00022507 | 0.17145
Bell Mass | M | 0.3 | kg | 5e-06 | kg | 2 | 1 | 5e-06 | 2.5e-06 | gl/Ae | 10001.5196 | 0.025004 | 30 | 0.00062519 | 0.28576
Local Gravity | gL | 9.8066 | m/sec2 | 0.0002 | % | 1 | 1 | 1.9613e-05 | 9.8066e-06 | gL-1 x P | 17845.0337 | 0.175 | 30 | 0.030625 | 2
Air Density | rAIR | 1.2 | kg/m3 | 0 | kg/m3 | 2 | 1 | 0 | 0 | rM-1 x P | 25 | 0 | 30 | 0 | 0
Ambient Pressure | PAMB | 97000 | Pa | 140 | Pa | 2 | 1 | 140 | 0 | 1.5E-9 x P | 0.2625 | 0 | 30 | 0 | 0
Ambient Temperature | TAMB | 23 | °C | 1 | °C | 2 | 1 | 1 | 0 | 5.5E-7 x P | 0.09625 | 0 | 30 | 0 | 0
Ambient Humidity | HAMB | 50 | %RH | 10 | %RH | 2 | 1 | 10 | 0 | 1.2E-8 x P | 0.0021 | 0 | 30 | 0 | 0
Mass Load AVG Density | rM | 7000 | kg/m3 | 40 | kg/m3 | 2 | 2 | 40 | 23.094 | rAIR/rM2 x P | 0.0042857 | 0.098974 | 30 | 0.0097959 | 1.1311
Surface Tension | τ | 0.032 | N/m | 10 | % | 1 | 2 | 0.0032 | 0 | πD/A0 | 113.2081 | 0 | 30 | 0 | 0
Verticality (<0.5 deg.) | q | 0 | deg. | 0.167 | deg. | 2 | 1 | 0.167 | 0.0835 | 1 - COSΦ x P | 0.18584 | 0.18584 | 30 | 0.034536 | 2.1239
System Terms
Density of Fluid | rFLUID | 2.1013 | kg/m3 | 0.1 | % | 1 | 1 | 0.0021013 | 0.0010506 | Aoh/M x P | 0.98067 | 0.0010303 | 30 | 1.0616e-06 | 0.011775
Fluid Head Height | h | 0.1 | m | 0.01 | m | 2 | 1 | 0.01 | 0.005 | g(rFLUID-rAIR) | 20.6063 | 0.10303 | 30 | 0.010616 | 1.1775
Float Position | fpi | 0.007 | m | 0.0001 | m | 2 | 1 | 0.0001 | 5e-05 | g(rFLUID-rAIR) | 20.6063 | 0.0010303 | 30 | 1.0616e-06 | 0.011775
Barometer | Pbaro | 97000 | Pa | 10 | Pa | 2 | 1 | 10 | 0 | 1 | 1 | 0 | 30 | 0 | 0
Reference Pressure | PREF | 4 | Pa | 10 | % | 1 | 1 | 0.4 | 0.2 | 1 | 1 | 0.2 | 30 | 0.04 | 2.2857
Additional | P | 175000 | Pa | 0 | Pa | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 30 | 0 | 0
[11 empty] | Combined Variance: | 0.98719 | Pa |  | Coverage
[11 empty] | Combined Uncertainty (k=1): | 0.99358 | Pa | k | 1.9955
[11 empty] | Expanded Uncertainty (95%): | 1.9826 | Pa | DOF | 68.598
[11 empty] | Expanded Uncertainty (95%): | 11.3294 | ppm

## Unccalc3 (hidden)
Limits
Pressure Value | 10000 | Pa | 1.4504 | psi | PABS | 10000
Mode | 1 | 1 for ABS, 2 for Gauge, 3 for ABS by ATM
Media :gas | 1 | 1 for gas, 2 for oil, 3 for water |  |  |  | Entry Field
KN (mass to pressure) | 10001.5196 | Pa/kg | 1.4506 | psi/kg |  | Not Used
Oil Density | 916 | kg/m3 |  |  |  | Password | Fluke
Maximum Head Height | 0.1 | m | 0.3937 | inches
Reference Temperature | 23 | °C What temperature is the effective area given
P-C Temperature Max Limit | 3 | °C What is the maximum difference in room temperature and reference temperature
Confidence | 95 | %
Influence Quantities |  |  |  | Influence Uncertainties | [5 empty] | Pressure Variant Terms
Influence | Typical | Approx. |  | Uncertainty Value |  | Type Value | Distribution | Uncertainty | Equiv. Std. | Sensitivity | Sensitivity | Std. Unc. | DOF | Variance | Expanded
Description | Symbol | Value | Units | Either % or value in unit 2σ or 95% |  | 1=relative(%), 2=constant | 1=normal, 2=rectangular |  | Unc. (1s) | Coefficient | Value | (Pa) | ----- | (Pa)² | (2σ, PPM)
Effective Area Terms
Effective Area | Ao | 0.00098052 | m2 | 0.0009 | % | 1 | 1 | 8.8246e-09 | 4.4123e-09 | Ao-1 x P | 10198711.6987 | 0.045 | 30 | 0.002025 | 9
Piston Thermal Coef. | aP | 4.55e-06 | °C-1 | 5 | % | 1 | 2 | 2.275e-07 | 1.3135e-07 | (TMAX - TREF) x P | 30000 | 0.0039404 | 30 | 1.5527e-05 | 0.78808
Cylinder Thermal Coef. | aC | 4.55e-06 | °C-1 | 5 | % | 1 | 2 | 2.275e-07 | 1.3135e-07 | (TMAX - TREF) x P | 30000 | 0.0039404 | 30 | 1.5527e-05 | 0.78808
Piston-Cylinder Temp. | T | 26 | °C | 0.1 | °C | 2 | 1 | 0.1 | 0.05 | (αp+αc) x P | 0.091 | 0.00455 | 30 | 2.0703e-05 | 0.91
1st Order Pressure Coef. | l,b1 | 4.2e-06 | MPa-1 | 0 | % | 1 | 1 | 0 | 0 | P2 | 100000000 | 0 | 30 | 0 | 0
2nd Order Pressure Coef. | b2 | 0 | MPa-2 | 0 | % | 1 | 1 | 0 | 0 | P3 | 1000000000000 | 0 | 30 | 0 | 0
Linearity | P | 10000 | Pa | 0.0002 | % | 1 | 1 | 0.02 | 0.01 | 1 | 1 | 0.01 | 30 | 0.0001 | 2
Performance | P | 10000 | Pa | 1.5e-05 | % | 1 | 1 | 0.0015 | 0.00075 | 1 | 1 | 0.00075 | 30 | 5.625e-07 | 0.15
Sensitivity Limit | P | 10000 | Pa | 0.015 | Pa | 2 | 1 | 0.015 | 0.0075 | 1 | 1 | 0.0075 | 30 | 5.625e-05 | 1.5
Ao Stability (2 yrs) | Ao | 0.00098052 | m2 | 0 | % | 1 | 2 | 0 | 0 | Ao-1 x P | 10198711.6987 | 0 | 30 | 0 | 0
Force Terms
Mass Load | M | 0.99985 | kg | 0.0005 | % | 1 | 1 | 4.9992e-06 | 2.4996e-06 | gl/Ae | 10001.5196 | 0.025 | 30 | 0.000625 | 5
Mass Stability | M | 0.99985 | kg | 0 | % | 1 | 2 | 0 | 0 | gl/Ae | 10001.5196 | 0 | 30 | 0 | 0
AMH Lubrication | M | 1.5e-05 | kg | 1e-05 | kg | 2 | 1 | 1e-05 | 5e-06 | gl/Ae | 10001.5196 | 0.050008 | 30 | 0.0025008 | 10.0015
Piston Mass | M | 0.2 | kg | 3e-06 | kg | 2 | 1 | 3e-06 | 1.5e-06 | gl/Ae | 10001.5196 | 0.015002 | 30 | 0.00022507 | 3.0005
Bell Mass | M | 0.3 | kg | 5e-06 | kg | 2 | 1 | 5e-06 | 2.5e-06 | gl/Ae | 10001.5196 | 0.025004 | 30 | 0.00062519 | 5.0008
Local Gravity | gL | 9.8066 | m/sec2 | 0.0002 | % | 1 | 1 | 1.9613e-05 | 9.8066e-06 | gL-1 x P | 1019.7162 | 0.01 | 30 | 0.0001 | 2
Air Density | rAIR | 1.2 | kg/m3 | 0 | kg/m3 | 2 | 1 | 0 | 0 | rM-1 x P | 1.4286 | 0 | 30 | 0 | 0
Ambient Pressure | PAMB | 97000 | Pa | 140 | Pa | 2 | 1 | 140 | 0 | 1.5E-9 x P | 0.015 | 0 | 30 | 0 | 0
Ambient Temperature | TAMB | 23 | °C | 1 | °C | 2 | 1 | 1 | 0 | 5.5E-7 x P | 0.0055 | 0 | 30 | 0 | 0
Ambient Humidity | HAMB | 50 | %RH | 10 | %RH | 2 | 1 | 10 | 0 | 1.2E-8 x P | 0.00012 | 0 | 30 | 0 | 0
Mass Load AVG Density | rM | 7000 | kg/m3 | 40 | kg/m3 | 2 | 2 | 40 | 23.094 | rAIR/rM2 x P | 0.0002449 | 0.0056557 | 30 | 3.1987e-05 | 1.1311
Surface Tension | τ | 0.032 | N/m | 10 | % | 1 | 2 | 0.0032 | 0 | πD/A0 | 113.2081 | 0 | 30 | 0 | 0
Verticality (<0.5 deg.) | q | 0 | deg. | 0.167 | deg. | 2 | 1 | 0.167 | 0.0835 | 1 - COSΦ x P | 0.010619 | 0.010619 | 30 | 0.00011277 | 2.1239
System Terms
Density of Fluid | rFLUID | 0.12007 | kg/m3 | 0.1 | % | 1 | 1 | 0.00012007 | 6.0036e-05 | Aoh/M x P | 0.98067 | 5.8875e-05 | 30 | 3.4663e-09 | 0.011775
Fluid Head Height | h | 0.1 | m | 0.01 | m | 2 | 1 | 0.01 | 0.005 | g(rFLUID-rAIR) | 1.1775 | 0.0058875 | 30 | 3.4663e-05 | 1.1775
Float Position | fpi | 0.007 | m | 0.0001 | m | 2 | 1 | 0.0001 | 5e-05 | g(rFLUID-rAIR) | 1.1775 | 5.8875e-05 | 30 | 3.4663e-09 | 0.011775
Barometer | Pbaro | 97000 | Pa | 10 | Pa | 2 | 1 | 10 | 0 | 1 | 1 | 0 | 30 | 0 | 0
Reference Pressure | PREF | 4 | Pa | 10 | % | 1 | 1 | 0.4 | 0.2 | 1 | 1 | 0.2 | 30 | 0.04 | 40
Additional | P | 10000 | Pa | 0 | Pa | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 30 | 0 | 0
[11 empty] | Combined Variance: | 0.046489 | Pa |  | Coverage
[11 empty] | Combined Uncertainty (k=1): | 0.21561 | Pa | k | 2.0211
[11 empty] | Expanded Uncertainty (95%): | 0.43577 | Pa | DOF | 40.2407
[11 empty] | Expanded Uncertainty (95%): | 43.577 | ppm

## Historical_Ae
Due Date | Aging Rate/ mm^2/day | Y intercept | Days Until Due Date | Number of days from last Cal | Number of Cals | Sum Square of Calibration Time | Sum of calibration time squared | Computation of Sxx | Mean calibration time | Standard Error of the Fit
2016-05-13 | -2.0729e-08 | 49.0199 | 219.6322 | 511.3678 | 9 | 81167002 | 467164996 | 29259780.2222 | 2401.5556 | 0.00016064
|  | Confidence based on todays date and next cal date | 0.00039477 | 0.00040163
Effective Area Calibration History |  |  | 8.0533 | 8.1932 | PPM
Date | Effective | Ae UNC | Ae UNC | Days | Pred. Value | Upper bound | Lwr bound
| Area (Ae) | 95% (k=2) | 95% (k=2)
| [mm2] | [ppm] | [mm2] |  | [mm2] | [mm2] | [mm2]
1999-11-02 | 49.0192 | 16 | 0.00078431 | 0 | 49.0192 | 49.0195 | 49.0188
2000-11-09 | 49.0194 | 16 | 0.00078431 | 373 | 49.0192 | 49.0195 | 49.0188
2001-11-12 | 49.0189 | 16 | 0.0007843 | 741 | 49.0192 | 49.0195 | 49.0188
2003-11-12 | 49.0191 | 10 | 0.00049019 | 1471 | 49.0191 | 49.0195 | 49.0188
2005-11-14 | 49.0192 | 10 | 0.00049019 | 2204 | 49.0191 | 49.0195 | 49.0188
2008-05-05 | 49.0189 | 8 | 0.00039215 | 3107 | 49.0191 | 49.0194 | 49.0188
2010-05-07 | 49.0191 | 8 | 0.00039215 | 3839 | 49.0191 | 49.0194 | 49.0187
2012-05-10 | 49.0191 | 8 | 0.00039215 | 4573 | 49.0191 | 49.0194 | 49.0187
2014-05-13 | 49.0192 | 8 | 0.00039215 | 5306 | 49.0191 | 49.0194 | 49.0187

## Historical_PCmass
Due Date | Aging Rate/ grams/day | Y intercept | Days Until Due Date | Number of days from last Cal | Number of Cals | Sum Square of Calibration Time | Sum of calibration time squared | Computation of Sxx | Mean calibration time | Standard Error of the Fit
2016-05-13 | -2.3822e-06 | 200.0808 | 219.6322 | 511.3678 | 12 | 108738399 | 912583681 | 32689758.9167 | 2517.4167 | 0.0016626
|  | Confidence based on todays date and next cal date | 0.0039575 | 0.004021
Effective Area Calibration History |  |  | 19.7796 | 20.0969 | PPM
Date | Effective |  | Mass UNC | Days | Pred. Value | Upper bound | Lwr bound
| Area (Ae) |  | 95% (k=2)
| [g] |  | [g] |  | [g] | [g] | [g]
1999-11-02 | 199.9972 |  | 0.003 | 0 | 199.9939 | 199.9977 | 199.9902
2000-11-09 | 199.994 |  | 0.003 | 373 | 199.993 | 199.9967 | 199.9893
2001-11-12 | 199.9919 |  | 0.003 | 741 | 199.9921 | 199.9958 | 199.9885
2003-11-12 | 199.989 |  | 0.003 | 1471 | 199.9904 | 199.9939 | 199.9869
2004-11-05 | 199.9876 |  | 0.003 | 1830 | 199.9896 | 199.993 | 199.9861
2005-11-14 | 199.9874 |  | 0.003 | 2204 | 199.9887 | 199.9921 | 199.9852
2006-11-09 | 199.987 |  | 0.003 | 2564 | 199.9878 | 199.9913 | 199.9843
2008-05-05 | 199.9853 |  | 0.003 | 3107 | 199.9865 | 199.99 | 199.983
2010-05-07 | 199.984 |  | 0.003 | 3839 | 199.9848 | 199.9883 | 199.9812
2011-05-04 | 199.9843 |  | 0.003 | 4201 | 199.9839 | 199.9875 | 199.9803
2012-05-10 | 199.9838 |  | 0.003 | 4573 | 199.983 | 199.9867 | 199.9794
2014-05-13 | 199.9835 |  | 0.003 | 5306 | 199.9813 | 199.9851 | 199.9775

## Revision
REVISION LOG: Piston Gauge Uncertainty Analysis
Revision Number | Date | Comments | Author | Function
1 | 20151006 | Initial creation of spreadsheet. | Michael Bair | Metrologist