Apply Offset and Slope correction to DUT, COMPASS Relationship Macro

'**************************************************

'This Function must return the device output

'in the Final Output Unit by manipulating the

'parameters provided.

'

'Out1     :By default, this is the raw formatted, output of a

'          device.  In other cases, it is final output of a

'          test configured device in the configured output unit.

'Out2     :Final output of a test configured device in the

'          configured output unit.

'Out3     :Final output of a test configured device in the

'          configured output unit.

'ParamID  :Parameter ID of the device .

'cCalc    :Calculation Class of the output.

'**************************************************

Function ApplyOffsetAndSlope(Out1, Out2, Out3, ParamID, cCalc)

'This proportional correction is used with DUTs with an offset and span correcton

'Use ApplyOffset for voltage DUTs and current 'DUTs with a simple offset correction

'This macro corrects with this formula

'Adjusted Final Pressure = (DUTRaw - Offset)* SlopeCorrection

n = cCOMPASS.cConfig.DUTPrs.Count

cyc = cCOMPASS.PressureCycle

cDebug.LogStatus "Pressure Cycle # " & cyc

For i = 1 to n

If cyc = 1 Then

Offset= GetVal(cCOMPASS.cConfig.DUTPrs(n).RangeMain.GetParent.CalInfo1)

'If Offset = 0 Then Exit Function

Slope=GetVal(ccompass.cConfig.DUTPrs(n).RangeMain.getparent.CalInfo2)

'If Slope=0 Then Exit Function

cDebug.LogStatus "Cycle 1, use CalCoef1 and CalCoef2 for offset and slope"

cDebug.LogStatus "SN: " & cCalc.RangeMain.GetParent.SN

cDebug.LogStatus "CalInfo1 (Offset): " & cCalc.RangeMain.GetParent.CalInfo1

cDebug.LogStatus "Offset: " & Offset

cDebug.LogStatus "CalInfo2 (Slope or Span): " & cCalc.RangeMain.GetParent.CalInfo2

cDebug.LogStatus "Slope (Span): " & Slope

ElseIf cyc = 2 Then

'	Offset= GetVal(cCOMPASS.cConfig.DUTPrs(n).RangeMain.GetParent.CalInfo4)

Offset = cCOMPASS.DataCollection(n).DUT.CalibrationCoefficient4

'If Offset = 0 Then Exit Function

'	Slope=GetVal(ccompass.cConfig.DUTPrs(n).RangeMain.getparent.CalInfo5)

Slope = cCOMPASS.DataCollection(n).DUT.CalibrationCoefficient5

'If Slope=0 Then Exit Function

cDebug.LogStatus "Cycle 2, use CalCoef4 and CalCoef5 for offset and slope"

cDebug.LogStatus "SN: " & cCalc.RangeMain.GetParent.SN

cDebug.LogStatus "CalInfo4 (Offset): " & cCalc.RangeMain.GetParent.CalInfo4

cDebug.LogStatus "Offset: " & Offset

cDebug.LogStatus "CalInfo5 (Slope or Span): " & cCalc.RangeMain.GetParent.CalInfo5

cDebug.LogStatus "Slope (Span): " & Slope

End If

MinRawOut = cCalc.RangeMain.MinRaw

cDebug.LogStatus "RawMin: " & MinRawOut

MaxRawOut = cCalc.RangeMain.MaxRaw

cDebug.LogStatus "RawMax: " & MaxRawOut

MinFinalOut = cCalc.RangeMain.MinFinal

cDebug.LogStatus "FinalMin:" & MinFinalOut

MaxFinalOut = cCalc.RangeMain.MaxFinal

cDebug.LogStatus "Final Max: " & MaxFinalOut

cDebug.LogStatus "Raw Output (Out1): " & Out1

Out1A = Out1 - Offset

cDebug.LogStatus "Adjusted Raw Output: " & Out1A

'Pct = (Out1A-MinRawOut)/(MaxRawOut-MinRawOut)

NewFinal = (Out1-Offset)*Slope

cDebug.LogStatus "Final Output:" & NewFinal

ApplyOffsetAndSlope = NewFinal

Next

End Function