## Sheet1
This data is the result of setting the E-DWT to 5000 psi and noting the displayed pressure value every 30 seconds.
You can see that after 2 minutes the apparent leak rate has dropped to less than 10% of the E-DWT measurement uncertainty specification.
Leaks and/or noise that are 10% or less of the desired measurement are generally considered to be insignificant.
And remember that the measurement uncertainty specification of the test device (UUT) is considerably worse than that of the E-DWT.
Pressure, psi | Time, minutes | Difference from the previous measurement, psi | Apparent Leak rate, psi/sec | Apparent Leak rate, ppm/sec | Apparent Leak rate, %FS/sec | E-DWT spec, ppm | [(E-DWT spec)/10], ppm
5000 | 0
4971.18 | 0.5 | -28.82 | -0.96067 | -192.1333 | -0.019213 | 200 | 20 | -20
4959.46 | 1 | -11.72 | -0.39067 | -78.5863 | -0.0078133 | 200 | 20 | -20
4955.14 | 1.5 | -4.32 | -0.144 | -29.0354 | -0.00288 | 200 | 20 | -20
4952.51 | 2 | -2.63 | -0.087667 | -17.6921 | -0.0017533 | 200 | 20 | -20
4951.4 | 2.5 | -1.11 | -0.037 | -7.471 | -0.00074 | 200 | 20 | -20
4950.07 | 3 | -1.33 | -0.044333 | -8.9537 | -0.00088667 | 200 | 20 | -20
4948.33 | 3.5 | -1.74 | -0.058 | -11.717 | -0.00116 | 200 | 20 | -20
4947.22 | 4 | -1.11 | -0.037 | -7.4773 | -0.00074 | 200 | 20 | -20
4946.01 | 4.5 | -1.21 | -0.040333 | -8.1527 | -0.00080667 | 200 | 20 | -20
4945.92 | 5 | -0.09 | -0.003 | -0.60655 | -6e-05 | 200 | 20 | -20
4946.26 | 5.5 | 0.34 | 0.011333 | 2.2915 | 0.00022667 | 200 | 20 | -20
4945.69 | 6 | -0.57 | -0.019 | -3.8413 | -0.00038 | 200 | 20 | -20
4945.03 | 6.5 | -0.66 | -0.022 | -4.4483 | -0.00044 | 200 | 20 | -20
4943.78 | 7 | -1.25 | -0.041667 | -8.426 | -0.00083333 | 200 | 20 | -20
4943.58 | 7.5 | -0.2 | -0.0066667 | -1.3485 | -0.00013333 | 200 | 20 | -20
4942.54 | 8 | -1.04 | -0.034667 | -7.0125 | -0.00069333 | 200 | 20 | -20
4942.6 | 8.5 | 0.06 | 0.002 | 0.40465 | 4e-05 | 200 | 20 | -20
4942.24 | 9 | -0.36 | -0.012 | -2.4279 | -0.00024 | 200 | 20 | -20
4942.15 | 9.5 | -0.09 | -0.003 | -0.60701 | -6e-05 | 200 | 20 | -20
4941.41 | 10 | -0.74 | -0.024667 | -4.9911 | -0.00049333 | 200 | 20 | -20