COMPASS for Pressure (or Flow)

Interface Macro to read ambient pressure, temperature and RH from a LEMCal log file

Copy and Paste all of the following text to an Interface Macro in the COMPASS Macro Editor. See the end of this document for screen-shots of the LEM setup in COMPASS for Pressure.

The macro is written for LEM serial number 57159. It will be updated in the future to read the serial number from COMPASS and also to automatically read the correct file. At present it requires the user to specify the current year, month and week numbers (need to edit the macro below appropriately).

The macro can be edited but below it is written for weekly log files written to this location,

C:\dhi\COMPASS for Pressure\Data\2456-LEM logs\SN057159_Y2016_W31.LOG

In LEMCal choose this location or edit the macro code below. The macro cheks that the last logged data value was today and less than 5 minutes ago. This can be edited if desired in the macro.

'**************************************************

'This Function must read and return the output of a remote instrument.

'The output must be in the Raw Output Unit defined by the Range class.

'

'cCalc     :The calculation class of the device.

'cParent   :The parent device class.

'fTarget1  :Target output for controller.

'bChnlSt   :True/False to activate or de-active a valve driver.

'cConfig   :Configuration class that holds all active calculation classes.

'

'The value is returned by setting the function name =

'to the calculated value..

'

'For example: Interface40270 =  val(mid(rawReply,5))

'**************************************************

'This macro reads from a log file created by LEMcal

'

'LEM Data can be logged into a Day, Week, Or Monthly file.

'Timing And Date information are based upon the clock of the host computer.

'	Day: Readings are taken at 1 minute intervals.

'		The active file Is saved And a New file Is created at midnight of each day.

'	Weekly: Readings are taken at 5 minute intervals.

'		The active file Is saved And a New file Is created at midnight of each Saturday.

'	Monthly: Readings are taken at 15 minute intervals.

'		The active file Is saved And a New file Is created at midnight of the last day of each month.

'

'Default Location

'By default, the log files are placed in

'C:\Program Files\Ruska Instrument\LEM Cal\LEMDataLog\

'or for 64-bit operating systems

'C:\Program Files (x86)\Ruska Instrument\LEM Cal\LEMDataLog\

'

'File names are generated automatically using the following formats

'	Day: SerialNumber_Year_Day of Year.LOG

'		Example: SN069362_Y2015_D137.LOG

'	Week: SerialNumber_Year_Week of Year.LOG

'		Example: SN069362_Y2015_W47.LOG

'	Month: SerialNumber_Year_Month of Year.LOG

'		Example: SN069362_Y2015_M05.LOG

'

'Data lines in the file are like this, units can be different, no header line in file

'7/22/2016 16:09:31  28.77 °C  24.15 %  97.36 kPa 1.119E-03 g/cm³ 000 chk 26577

Function ReadLEMCal_logFile(cCalc, cParent, fTarget1, bChnlSt, cConfig)

Dim fso, f

YYYY = 2016

MM = 07

WW=31

'fName = "C:\dhi\COMPASS for Pressure\Data\2456-LEM logs\SN057159_Y2016_W31.LOG"

fName = "C:\dhi\COMPASS for Pressure\Data\2456-LEM logs\SN057159_Y" & YYYY & "_W" & WW & ".LOG"

'fName = "C:\Program Files (x86)\Ruska Instrument\LEM Cal\LEMDataLog\SN069362_Y2016_M07.LOG"

'fName = "C:\Program Files\Ruska Instrument\LEM Cal\LEMDataLog\"

'fName = "C:\dhi\COMPASS for Pressure\Data\2456-LEM logs\LEM log file.log"

cDebug.LogStatus "File Name = " & fName

Set fso = CreateObject("Scripting.FileSystemObject")

Set f = fso.OpenTextFile(fName, 1)

fData = f.ReadAll   'Read the entire file into that variable

allData = Split(fData, vbcrlf)

i = Ubound(allData)

cDebug.LogStatus "line " & (i) & " in log file = " & allData(i-1)

'	cDebug.LogStatus "line " & "1" & " in log file = " & allData(1)

'	cDebug.LogStatus "line " & (i-2) & " in log file = " & allData(i-2)

LastLine = allData(i-1)

'	cDebug.LogStatus "just data = " & LastLine

'Example line in log file

'7/25/2016 17:45:00  23.54 °C  44.55 %  96.50 kPa 1.127E-03 g/cm³ 000 chk 57855

LineData = Split(LastLine," ")	'split the last line of the log file into components separated by spaces

Dt = CDate(Trim(LineData(0)))

'	cdebug.LogStatus "Date = " & Dt

Tm = CDate(LineData(1))

'	cdebug.LogStatus "Time = " & Tm

DtTm = CDate(Dt & " " & Tm)

'	cDebug.LogStatus "Date & Time = " & DtTm

'	cDebug.LogStatus "Date & Time Now = " & Date & " " & Time

'Check that time is within 5 minutes of now

'	cDebug.LogStatus "DateDiff = " & DateDiff(dateinterval.weekday,dt,Date)

DateDelta = Date - Dt	'In days

TimeDelta = formatNumber(((Time - Tm) * 24 * 60),2)	'Convert days to minutes

cDebug.LogStatus "DateDelta = " & DateDelta & " days" & ", TimeDelta = " & TimeDelta & " minutes"

If abs(DateDelta) > 0 Then

msgbox "Last logged LEM data is from " & Dt & ", (not today) ending test"

cCOMPASS.SystemAbort = True

TimeDelay 1

Exit Function

End If

If abs(TimeDelta)> 5 Then

msgbox "Last logged LEM data is from " & DtTm & ", and is more than 5 minutes ago, ending test"

cCOMPASS.SystemAbort = True

'	TimeDelay 1

Exit Function

End If

PrsUtestRaw = cConfig.AmbPrs.RangeMain.UnitRaw

cDebug.LogStatus "PrsUtestRaw = " & PrsUtestRaw	'mbar=2, kPa=3,psi=9, hPa=38

PrsUtestFnl = cConfig.AmbPrs.RangeMain.UnitFinal

cDebug.LogStatus "PrsUtestFinal = " & PrsUtestFnl

TempUtest = cConfig.AmbTemp.RangeMain.UnitFinal

cDebug.LogStatus "TempUtest = " & TempUtest	'C=0, F=1, K=2, R=3

HumUtest = cConfig.AmbRH.RangeMain.UnitFinal

cDebug.LogStatus "HumUtest = " & HumUtest

'Example data line and where the splits are based on the space character

'0         1        23     4  56     7 89     10  11        12    13  14  15

'7/25/2016 17:45:00  23.54 °C  44.55 %  96.50 kPa 1.127E-03 g/cm³ 000 chk 57855

'	cDebug.LogStatus "0 = " & LineData(0)

'	cDebug.LogStatus "1 = " & LineData(1)

'	cDebug.LogStatus "2 = " & LineData(2)

'	cDebug.LogStatus "3 = " & LineData(3)

'	cDebug.LogStatus "4 = " & LineData(4)

'	cDebug.LogStatus "5 = " & LineData(5)

'	cDebug.LogStatus "6 = " & LineData(6)

'	cDebug.LogStatus "7 = " & LineData(7)

'	cDebug.LogStatus "8 = " & LineData(8)

'	cDebug.LogStatus "9 = " & LineData(9)

'	cDebug.LogStatus "10 = " & LineData(10)

PrsAmbRaw = LineData(9)

PrsUraw = LineData(10)

cdebug.LogStatus "P_amb from file = " & PrsAmbRaw & " " & PrsUraw

'PrsUnumRaw = 3	'kPa

'Reference Macro Editor Help for details on the UnitConversion function

'cCOMPASS.UnitConversion dInput, iNewUnit, iCurUnit, iType

'PrsAmb = cCOMPASS.UnitConversion(CDbl(PrsAmbRaw),CInt(PrsUtestRaw),CInt(PrsUnumRaw),0) 'always 0 for pressure devices, 10 for RH

PrsAmb = PrsAmbRaw

cdebug.LogStatus "Prs_amb = " & PrsAmb & " " & cConfig.AmbPrs.RangeMain.UnitFinalText

TempAmbRaw = LineData(3)

TempUraw = Right(LineData(4),1)

cdebug.LogStatus "T_amb from file = " & TempAmbRaw & " " & TempUraw

TempUnumRaw = 0	'deg Celsius

TempAmb = cCOMPASS.UnitConversion(CDbl(TempAmbRaw),CInt(TempUtest),CInt(TempUnumRaw),1) 'always 1 for temp. devices

cdebug.LogStatus "Temp_amb = " & TempAmb & " " & cConfig.AmbTemp.RangeMain.UnitFinalText

HumAmb = LineData(6)

HumUraw = LineData(7)

cdebug.LogStatus "RH_amb = " & HumAmb & " " & HumUraw

cCOMPASS.cConfig.AmbPrs.RangeMain.RawOutput = PrsAmb

cCOMPASS.cConfig.AmbTemp.RangeMain.RawOutput = TempAmb

cCOMPASS.cConfig.AmbRH.RangeMain.RawOutput = HumAmb

End Function