SECTION 4.0

STARTUP AND CONFIGURATION

4.1	CONNECTING EXTERNAL COMPONENTS

Prior to startup the following components must be properly connected to the AutoPrompter:

1.	Standard 115 VAC, 60 Hz power.

2.	RTD Sensor.

3.	FPI Sensor (optional; recommended).

4.	Printer (optional).

Refer to Figure 4-1 for an illustration of the AutoPrompter's back panel.

Refer to Appendix C for instructions on connecting RTD and FPI sensors to the deadweight gage.

AutoPrompt Converter Back Panel

Figure 4-1

4.1.1	CONNECTING AN RTD SENSOR

The RTD sensor is mounted in the base of the deadweight gage.  The sensor is connected to the AutoPrompter by a cable.  This cable connects to the AutoPrompter back panel at the location labelled 'RTD'.

The RTD sensor is optional but recommended.

4.1.2	CONNECTING AN FPI SENSOR

The FPI sensor cable connects to the AutoPrompter's back panel via a BNC connector labeled 'FPI'.  The FPI sensor is mounted in the base of the deadweight gage.  The FPI sensor is optional but recommended.

The FPI sensor is optional but recommended.

4.1.3	CONNECTING A PRINTER

An RS-232 printer interfaces to the AutoPrompter at the back panel DB-25 connector labeled 'PRINTER'.  The printer and cable are normally not supplied by the factory.

4.2	STARTUP

4.2.1	POWERING ON

The user should select either 115 VAC or 230 VAC; 50 or 60 Hz by proper positioning of the voltage selector on the back of the AutoPrompter.

With the power cord disconnected to both the AutoPrompter and wall outlet, use a small screwdriver to pry open the voltage select/fuse compartment on the power inlet module on the back of the AutoPrompter.  Install the correct fuses for local power supply conditions in their receptacles.  Remove the voltage select wheel and turn it so that when it is replaced within its compartment and the cover is closed, the correct voltage is displayed through the window.

Finally, connect the power and power up the AutoPrompter by flipping the power on/off switch located on the back panel.

CAUTION:  DO NOT attempt to turn the voltage select wheel while it is still in place within the fuse compartment.  Always remove it first from the compartment and then replace it.  Turning the wheel while in place may damage the electrical contacts.

CAUTION: The AutoPrompter requires air circulation through vents located on the bottom and top.  Do not allow these vents to become blocked.

4.2.2	SIGN-ON DISPLAY

The following sign-on messages are displayed on the key panel at power up:

After about two seconds, the following screen is displayed:

where:

rRvv = software release and version number

sssss = 5-digit software set number

mm-dd-yy = month-day-year

hr:mn = hour:minute

The AutoPrompter will beep three times indicating proper sign-on.

After about three or four seconds, Main Menu #1 is displayed.

4.2.3	SELF-TEST

At power-up, the start-up sequence tests internal non-volatile memory structures for bad data.  If bad data is detected, one of the following occurs:

1.	An attempt is made to recover the affected memory structure(s) to a previously known state or;

2.	The affected memory structure is initialized to factory default values if it cannot be recovered.

If the memory structure can be recovered, then the following message is displayed:

Code 'x' identifies the affected memory structure and is discussed below in detail.

The user should check the data in the affected memory structure and edit as necessary.

If the affected memory structure is initialized to factory default values, then the following message is displayed:

The code 'x' indicates which memory structure has been affected.

These codes are identified as:

CODE		TYPE OF STRUCTURE		         SECTIONS TO REFERENCE

1	System Data	5.5, 5.7, 5.8, 5.11

2	FPI Sensor Adjustment Data	5.10.3

3	RTD Sensor Adjustment and Calibration Data	5.10.1, 5.10.2

4	Environment Parameter Data	5.6.3, 5.9.3

5	Float Position Parameter Data	5.11.3, 5.11.4

6	Piston/Cylinder Data Set No. 1	5.6.1, 5.9.1

7	Piston/Cylinder Data Set No. 2	5.6.1, 5.9.1

8	Piston/Cylinder Data Set No. 3	5.6.1, 5.9.1

9	Piston/Cylinder Data Set No. 4	5.6.1, 5.9.1

10	Piston/Cylinder Data Set No. 5	5.6.1, 5.9.1

11	Piston/Cylinder Data Set No. 6	5.6.1, 5.9.1

12	Mass Data Set No. 1	5.6.2, 5.9.2

13	Mass Data Set No. 2	5.6.2, 5.9.2

System data is affected by code 1 and includes:

1.	Printer configuration parameters (default is NOT INSTALLED, 9600 baud, no parity, 8 data bits, 1 stop bit, form feed ENABLED).

2.	Active piston/cylinder data set (default is 1).

3.	Active mass data set (default is 1).

4.	Active pressure unit (default is psi).

5.	RTD sensor installation option (default is NOT INSTALLED).

6.	FPI sensor installation option (default is NOT INSTALLED).

7.	FPI digital filter selection (none, low-pass or Butterworth; default is NONE).

FPI sensor adjustment data is affected by code 2 and includes:

1.	Float position adjustment coefficients.

2.	Last adjustment date and time.

RTD sensor adjustment and calibration data are affected by code 3 and include:

1.	'Auto-cal' adjustment coefficients.

2.	Offset calibration coefficient.

3.	Last calibration date and time for the RTD sensor offset.

4.	Last 'auto-cal' date and time.

Environment parameters are affected by code 4 and include:

1.	Active reference pressure (default is GAGE).

2.	Head Unit (default is CM).

3.	Oil head (default is 0.000 cm).

4.	Gas head (default is 0.00 cm).

5.	User-specified air density (default is 0.00000 g/cm3).

6.	User-specified local gravity (default is 0.0000 cm/sec2).

7.	Local elevation (for formula-based gravity calculation).

8.	Local latitude (for formula-based gravity calculation).

9.	Local air pressure (for formula-based air density calculation).

10.	Local temperature (for formula-based air density calculation).

11.	Local humidity (for formula-based air density calculation).

12.	Type of gas medium (none, Nitrogen, Helium or zero air; default is NONE).

13.	Type of oil medium (none, Spinesstic-22, DOS Sebacate; default is NONE).

14.	Active float position units (default is centimeter).

15.	Active graphic display window limit for float position (default is 0.381 cm).

Float position parameters are affected by code 5 and include:

1.	'A' and 'B' matrix coefficients for Butterworth filter (required by Ruska 2485 deadweight gage and similar gages).

Piston/cylinder data sets 1 through 6 are affected by codes 6 through 11, respectively.  Each set includes:

1.	Model number.

2.	Serial number.

3.	Medium (oil or gas, default is OIL).

4.	Tare weight and unit (kg or lbm) in both apparent mass and true mass.

5.	Ao, b1, b2, c and L dimension coefficients.

6.	Coefficient units (SI or ENGLISH; default is SI).

7.	Piston temperature reference (default is 23.0 degrees C).

8.	Piston tare mass'apparent mass standard reference (default is 8.4 g/cm3 for Brass).

Mass data sets 1 and 2 are affected by codes 12 and 13, respectively, and include:

1.	Model number.

2.	Serial number.

3.	Number of masses in the set.

4.	Mass units (kg or lbm, default is KG).

5.	Apparent Mass standard reference (default is 8.4 g/cm3 for Brass).

6.	Weight of each mass in both apparent mass and true mass.

7.	D Dimension coefficient in units of m or in; (default in m).

8.	Mass set break point for identifying mass dependencies.

The user should acknowledge each message by pressing the ENTER key.

CAUTION:  The user must take action to re-establish appropriate coefficients and data in memory.  If an error is detected, the user must identify the structure affected and must re-enter (or at least verify) the affected data.  If an error occurs in the FPI or RTD sensor data structures, then recalibration of those sensors is mandatory.

NOTE:  The very first time that the AutoPrompter is powered up (usually at the factory), all memory structures will report an error.  This is normal and establishes default values for the structures.

4.3	INITIAL CONFIGURATION OF PARAMETERS

Initially, the following parameters must be configured by the user:

1.	SYSTEM PARAMETERS; Sections 5.5, 5.7, 5.8, 5.11

a.	FPI sensor installation option; Section 5.11; factory default is NOT INSTALLED, filter selection of NONE.

b.	RTD sensor installation option; Section 5.11; factor default is NOT INSTALLED.

c.	Printer option; Section 5.11; factory default is NOT INSTALLED

d.	Printer interface parameters (if installed); Section 5.11; factory defaults are 9600 baud, no parity, 8 data bits, 1 stop bit; form feed disabled.

e.	Active piston/cylinder data set; Section 5.7; factory default is 1.

f.	Active mass data set; Section 5.8; factory default is 1.

g.	Current pressure units; Section 5.5; factory default is 'psi'.

2.	SENSOR ADJUSTMENT AND CALIBRATION PARAMETERS; Section 5.10

a.	RTD sensor adjustment and calibration coefficients

b.	FPI sensor adjustment coefficients

3.	PISTON/CYLINDER SETS' DATA AND COEFFICIENTS (for up to 6 sets); Section 5.9.1

4.	MASS SETS' DATA (for up to 2 mass sets); Section 5.9.2

5.	ENVIRONMENT PARAMETERS; Section 5.9.3

a.	Reference pressure of deadweight gage; gage,or absolute mode; factory default is GAGE.

b.	Pressure head unit and values; factory default head unit is CM.

c.	Local gravity at the reference plane of the deadweight gage

d.	Local air density at the reference plane of the deadweight gage

e.	Gas medium selection; None Nitrogen, Helium or Zero Air; factory default is NONE.

f.	Oil Medium selection; None, Spinnestic 22 or DOS Sebacate; factory default is NONE.

g.	Float position unit and graphic window limits.  Defaults are CM and 0.381 cm (0.150 in).

6.	DATE AND TIME; Section 5.9.5

The system parameters are configured from the '9-Configure System' selection of Main Menu #2.  Refer to Section 5.11 for a complete description on configuring these parameters.

The calibration/adjustment parameters are established by performing calibration/adjustment sequences.  These sequences are initiated from the '8-Calibrate Sensors' selection of Main Menu #2.  Refer to Section 5.10 for additional help.

The piston/cylinder sets' data and mass sets' data are established by user data entry.  Refer to Sections 5.9.1 and 5.9.2.

The environment parameters are configured by selecting the '7-Set Parameters' selection of Main Menu #2 and then selecting the '3-Set Environment Parameters' choice.  Refer to Section 5.9.3 (Set Environment Parameters).

The date and time are configured by selecting the '7-Set Parameters' selection of Main Menu #2 and then selecting the '4-Set Date/Time' choice.  Refer to Section 5.9.4 (Set Date and Time).

4.4	TURNING OFF THE AUTOPROMPTER

The AutoPrompter can be safely powered down while in Main Menu #1 or Main Menu #2.

CAUTION:  Powering down the AutoPrompter while entering or editing environment parameters, float position parameters, configuration parameters, mass set data or piston/cylinder set data can result in computer memory data loss.  To ensure data integrity of various memory structures, the AutoPrompter should only be powered down in Main Menu #1 or Main Menu #2.  Quick power down and power up cycles should be avoided.

THIS PAGE INTENTIONALLY LEFT BLANK