## Q-RPT
See PMM sheet for example PM600 module
PPC4-ui A700Kp/A200Kp S/N 964 |  |  | Calibration | AutoZ | AutoZ | AutoZ | Calibration | AutoZ | AutoZ | AutoZ | Calibration
|  |  | 2011-01-01 | 2011-04-01 | 2011-07-01 | 2011-10-01 | 2012-01-01 | 2012-04-01 | 2012-07-01 | 2012-10-01 | 2013-01-01
A700Kp | 0.008% | PA value (Pascal) | 0 | 0 | 0 | 0 | 2.2 | 2.2 | 2.2 | 2.2 | 1.7
FS, kPa | 700 | PM value (multiplier, no unit) | bar: 2011-01-01 to 2013-01-01
1-yr drift, Pa | 43.106 | AutoZero value (Pascal) | 0 | 2.1 | 3.2 | 2 | 0 | -0.6 | -0.4 | -0.7 | 0
|  | AutoZero + PA | 0 | 2.1 | 3.2 | 2 | 2.2 | 1.6 | 1.8 | 1.5 | 1.7
A200Kp | 0.008% | PA value (Pascal) | 0 | 0 | 0 | 0 | -0.4 | -0.4 | -0.4 | -0.4 | -0.5
FS, kPa | 200 | PM value (multiplier, no unit) | 1 | 1 | 1 | 1 | 0.99999 | 0.99999 | 0.99999 | 0.99999 | 0.99999
1-yr drift, Pa | 18.106 | AutoZero value (Pascal) | 0 | -0.4 | -0.3 | -0.5 | 0 | -0.1 | -0.1 | -0.2 | 0
|  | AutoZero + PA | 0 | -0.4 | -0.3 | -0.5 | -0.4 | -0.5 | -0.5 | -0.6 | -0.5
RPM4 A700Ks/A200Ks S/N 1106 |  |  | Calibration | AutoZ | AutoZ | AutoZ | Calibration | AutoZ | AutoZ | AutoZ | Calibration
|  |  | 2011-01-01 | 2011-02-01 | 2011-07-01 | 2011-10-01 | 2012-01-01 | 2012-04-01 | 2012-07-01 | 2012-10-01 | 2013-01-01
A700Ks | 0.01% | PA value (Pascal) | 0 | 0 | 0 | 0 | -0.7 | -0.7 | -0.7 | -0.7 | -0.8
FS, kPa | 700 | PM value (multiplier, no unit) | bar: 2011-01-01 to 2013-01-01
1-yr drift, Pa | 56 | AutoZero value (Pascal) | 0 | -1.2 | -0.6 | -0.8 | 0 | -0.2 | -0.3 | -0.2 | 0
|  | AutoZero + PA | 0 | -1.2 | -0.6 | -0.8 | -0.7 | -0.9 | -1 | -0.9 | -0.8
A200Kp | 0.01% | PA value (Pascal) | 0 | 0 | 0 | 0 | 0.9 | 0.9 | 0.9 | 0.9 | 0.8
FS, kPa | 200 | PM value (multiplier, no unit) | 1 | 1 | 1 | 1 | 0.99998 | 0.99998 | 0.99998 | 0.99998 | 0.99999
1-yr drift, Pa | 20.1325 | AutoZero value (Pascal) | 0 | 1 | 0.8 | 1.1 | 0 | -0.2 | -0.3 | -0.1 | 0
|  | AutoZero + PA | 0 | 1 | 0.8 | 1.1 | 0.9 | 0.7 | 0.6 | 0.8 | 0.8
PPCH-70M A70MS/A20MS S/N 673 |  |  | Calibration | AutoZ | AutoZ | AutoZ | Calibration | AutoZ | AutoZ | AutoZ | Calibration
|  |  | 2011-01-01 | 2011-02-01 | 2011-07-01 | 2011-10-01 | 2012-01-01 | 2012-04-01 | 2012-07-01 | 2012-10-01 | 2013-01-01
A70Ms | 0.013% | PA value (Pascal) | 0 | 0 | 0 | 0 | 198 | 198 | 198 | 198 | 123
FS, kPa | 70000 | PM value (multiplier, no unit) | 1 | 1 | 1 | 1 | 0.99997 | 0.99997 | 0.99997 | 0.99997 | 1
1-yr drift, Pa | 6230 | AutoZero value (Pascal) | 0 | 185 | 197 | 201 | 0 | -50 | -47 | -76 | 0
|  | AutoZero + PA | 0 | 185 | 197 | 201 | 198 | 148 | 151 | 122 | 123
A20Ms | 0.013% | PA value (Pascal) | 0 | 0 | 0 | 0 | -87 | -87 | -87 | -87 | -140
FS, kPa | 20000 | PM value (multiplier, no unit) | bar: 2011-01-01 to 2013-01-01
1-yr drift, Pa | 1780 | AutoZero value (Pascal) | 0 | -94 | -87 | -93 | 0 | -11 | -14 | -76 | 0
|  | AutoZero + PA | 0 | -94 | -87 | -93 | -87 | -98 | -101 | -163 | -140
RPM4 BA100K S/N 1234 |  |  | Calibration | AutoZ | AutoZ | AutoZ | Calibration | AutoZ | AutoZ | AutoZ | Calibration
|  |  | 2011-01-01 | 2011-02-01 | 2011-07-01 | 2011-10-01 | 2012-01-01 | 2012-04-01 | 2012-07-01 | 2012-10-01 | 2013-01-01
BA100K | 0.01% | PA value (Pascal) | 0 | 0 | 0 | 0 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
FS, kPa | 110 | PM value (multiplier, no unit) | 1 | 1 | 1 | 1 | 0.99999 | 1 | 1 | 1 | 0.99999
1-yr drift, kPa | 15.6325 | AutoZero value (Pascal) | 0 | -0.33 | -0.1 | 0.1 | 0 | -0.33 | -0.1 | 0.1 | 0
|  | AutoZero + PA | 0 | -0.33 | -0.1 | 0.1 | 0.2 | -0.13 | 0.1 | 0.3 | 0.2
Meas. Unc. Spec. at atm
greater of [((%Rdg Spec)*101.325kPa) or ((%Rdg Spec)*Threshold%*FS)]
1-year drift spec. at atm
0.005%FS + measurement Unc. Spec. at atm
Threshold%
Standard | 30%
Premium | 9%

## PMM
See Q-RPT sheet for Q-RPTs
PM600-A700K |  |  |  | Calibration | AutoZ | AutoZ | AutoZ | Calibration | AutoZ | AutoZ | AutoZ | Calibration
|  |  |  | 2016-01-01 | 2016-04-01 | 2016-07-01 | 2016-10-01 | 2017-01-01 | 2017-04-01 | 2017-07-01 | 2017-10-01 | 2018-01-01
%Rdg = | 0.01% |  | C0 value (psi) | 0.0012156 | 0.0012156 | 0.0012156 | 0.0012156 | 0.001487 | 0.001487 | 0.001487 | 0.001487 | 0.001397
FS, psi = | 100 |  | C1 value (multiplier, no unit) | bar: 2016-01-01 to 2018-01-01
1-yr drift, psi = | 0.0040462 |  | AutoZero value (psi), aka zOffset | 0 | 0.00030457 | 0.0004641 | 0.00029007 | 0 | -8.702e-05 | -5.8013e-05 | -0.00010152 | 0
|  |  | AutoZero + C0 | 0.0012156 | 0.0015202 | 0.0016797 | 0.0015057 | 0.001487 | 0.0014 | 0.001429 | 0.0013855 | 0.001397
PM600-A40M, Quartz Sensor |  |  |  | Calibration | AutoZ | AutoZ | AutoZ | Calibration | AutoZ | AutoZ | AutoZ | Calibration
|  |  |  | 2016-01-01 | 2016-04-01 | 2016-07-01 | 2016-10-01 | 2017-01-01 | 2017-04-01 | 2017-07-01 | 2017-10-01 | 2018-01-01
%Rdg = | 0.01% |  | C0 value (psi) | 0.057133 | 0.057133 | 0.057133 | 0.057133 | 0.059047 | 0.059047 | 0.059047 | 0.059047 | 0.005687
FS, psi = | 100 |  | C1 value (multiplier, no unit) | 0.99999 | 0.99999 | 0.99999 | 0.99999 | 0.99998 | 0.99998 | 0.99998 | 0.99998 | 0.99997
1-yr drift, psi = | N/A |  | AutoZero value (psi), aka zOffset | 0 | 0.001247 | 0.001892 | 0.001198 | 0 | -0.002547 | -0.002457 | -0.003547 | 0
|  |  | AutoZero + C0 | 0.057133 | 0.05838 | 0.059025 | 0.058331 | 0.059047 | 0.0565 | 0.05659 | 0.0555 | 0.005687
PM630-A40M, Onboard Barometer Sensor (used to automatically adjust AutoZero Quartz Sensor) |  |  |  | 2016-01-01: Calibration; 2017-01-01: Calibration; 2018-01-01: Calibration
|  |  |  | 2016-01-01 |  |  |  | 2017-01-01 |  |  |  | 2018-01-01
%FS | 0.38% |  | C0 value (psi) | 2016-01-01: -0.010105; 2017-01-01: 0.00031907; 2018-01-01: 0.00024656
FS, psi = | 16 |  | C1 value (multiplier, no unit) | 2016-01-01: 1.0026; 2017-01-01: 1; 2018-01-01: 1
1-yr drift, psi = | N/A
PM630-A100M, Quartz Sensor |  |  |  | Calibration | AutoZ | AutoZ | AutoZ | Calibration | AutoZ | AutoZ | AutoZ | Calibration
|  |  |  | 2016-01-01 | 2016-04-01 | 2016-07-01 | 2016-10-01 | 2017-01-01 | 2017-04-01 | 2017-07-01 | 2017-10-01 | 2018-01-01
%Rdg = | 0.012% |  | C0 value (psi) | 0.1714 | 0.1714 | 0.1714 | 0.1714 | 0.7026 | 0.7026 | 0.7026 | 0.7026 | 0.54785
FS, psi = | 100 |  | C1 value (multiplier, no unit) | 1 | 1 | 1 | 1 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | 1
1-yr drift, psi = | N/A |  | AutoZero value (psi), aka zOffset | 0 | 0.3215 | 0.42505 | 0.52141 | 0 | -0.2087 | -0.32146 | -0.25419 | 0
|  |  | AutoZero + C0 | 0.1714 | 0.4929 | 0.59645 | 0.69281 | 0.7026 | 0.4939 | 0.38115 | 0.44842 | 0.54785
PM630-A100M, Onboard Barometer Sensor (used to automatically adjust AutoZero Quartz Sensor) |  |  |  | 2016-01-01: Calibration; 2017-01-01: Calibration; 2018-01-01: Calibration
|  |  |  | 2016-01-01 |  |  |  | 2017-01-01 |  |  |  | 2018-01-01
%FS | 0.38% |  | C0 value (psi) | 2016-01-01: 0.020065; 2017-01-01: 0.00031907; 2018-01-01: 0.00024656
FS, psi = | 16 |  | C1 value (multiplier, no unit) | 2016-01-01: 1.0026; 2017-01-01: 1; 2018-01-01: 1
1-yr drift, psi = | N/A
Quartz sensor 1-year drift spec. at atm (absolute mode stability), k=2
0.007% of span
Quartz Sensor Threshold for % of Reading Specification
Threshold% = | 30% | All modules except PM630-A100M
Threshold% = | 40% | PM630-A100M
Measurement Uncertainty for A20M and below with 1-year drift (absolute mode stability) spec.
Onboard barometer 1-year specification
0.38% FS | measuremenent uncertainty
0.027% | precision