## Unccalc1
Limits | [8 empty] | System Uncertainty | [6 empty] | 1 | 1
Max Pressure Value | 1750000 | Pa | 253.816 | psi | PABS | 1750000 | Enter Min P (Pa) | 25000 |  | Pressure (Pa) | U(95) (Pa) | U(95) (PPM) |  |  |  | 2 | 2
Mode | 1 | 1 for ABS, 2 for Gauge, 3 for ABS by ATM, 4 DIFF mode |  |  |  |  | psi | 3.6259 | Max | 1750000 | 22.2671 | 12.7241 | Absolute by application of vacuum mode |  |  | 3 | 3
Media :gas | 1 | 1 for gas, 2 for oil, 3 for water (must be 1 if mode is 1) |  |  |  |  | key |  | Mid | 875000 | 11.1325 | 12.7229 |  |  |  | 4
KN (mass to pressure) | 50000.7648 | Pa/kg | 7.252 | psi/kg |  | Entry Field |  |  | Min | 25000 | 0.47289 | 18.9156
Oil Density | 916 | kg/m3 |  |  |  | Not Used
Maximum Head Height | 0.1 | m | 0.3937 | inches |  | Password | Fluke |  | Model/SN | Example PC-7102/7601-50
Reference Temperature | 20 | °C What temperature is the effective area given | [5 empty] | psi to Pa conv. |  | 0.0013% of rdg | + 0.13 Pa
P-C Temperature Max Limit | 5 | °C What is the maximum difference in p-c temperature and reference temperature | [5 empty] | 1
Confidence | 95 | % | [5 empty] | 6894.7591
Influence Quantities |  |  |  | Influence Uncertainties | [5 empty] | Pressure Variant Terms
Influence | Typical | Approx. |  | Uncertainty Value |  | Type Value | Distribution | Uncertainty | Equiv. Std. | Sensitivity | Sensitivity | Std. Unc. | DOF | Variance | Expanded
Description | Symbol | Value | Units | Either % or value in unit 2σ or 95% |  | 1=relative(%), 2=constant | 1=normal, 2=rectangular |  | Unc. (1s) | Coefficient | Value | (Pa) | ----- | (Pa)² | (2σ, PPM)
Effective Area Terms
Effective Area (B12) | AE(Tref,0) | 0.00019613 | m2 | 0.0011 | % | 1 | 1 | 2.1574e-09 | 1.0787e-09 | Ao-1 x P | 8922653342.171 | 9.625 | 30 | 92.6406 | 11
Piston Thermal Coef. (B16) | aP | 4.5e-06 | °C-1 | 2 | % | 1 | 2 | 9e-08 | 5.1962e-08 | (TMAX - TREF) x P | 8750000 | 0.45466 | 30 | 0.20672 | 0.51962
Cylinder Thermal Coef. (B16) | aC | 4.5e-06 | °C-1 | 2 | % | 1 | 2 | 9e-08 | 5.1962e-08 | (TMAX - TREF) x P | 8750000 | 0.45466 | 30 | 0.20672 | 0.51962
Piston-Cylinder Temp. (B10) | T | 25 | °C | 0.1 | °C | 2 | 1 | 0.1 | 0.05 | (αp+αc) x P | 15.75 | 0.7875 | 30 | 0.62016 | 0.9
1st Order Pressure Coef. (B15) | l,b1 | -2.6e-06 | MPa-1 | 10 | % | 1 | 1 | -2.6e-07 | -1.3e-07 | P2 | 3062500000000 | -0.39813 | 30 | 0.1585 | 0.455
2nd Order Pressure Coef. | b2 | 0 | MPa-2 | 0 | % | 1 | 1 | 0 | 0 | P3 | 5.3594e+18 | 0 | 30 | 0 | 0
Linearity (B14) | P | 1750000 | Pa | 0.00015 | % | 1 | 1 | 2.625 | 1.3125 | 1 | 1 | 1.3125 | 30 | 1.7227 | 1.5
Sensitivity/Performance (B13) | P | 1750000 | Pa | 1.5e-05 | % | 1 | 1 | 0.2625 | 0.13125 | 1 | 1 | 0.13125 | 30 | 0.017227 | 0.15
Sensitivity Limit (B13) | P | 1750000 | Pa | 0.058 | Pa | 2 | 1 | 0.058 | 0.029 | 1 | 1 | 0.029 | 30 | 0.000841 | 0.033143
Ao Stability (B17) | Ao | 0.00019613 | m2 | -4.6305e-05 | % | 1 | 2 | -9.0818e-11 | -5.2434e-11 | Ao-1 x P | 8922653342.171 | -0.46785 | 30 | 0.21888 | 0.53468
Force Terms
Mass Load (B1) | M | 34.9995 | kg | 0.0005 | % | 1 | 1 | 0.000175 | 8.7499e-05 | gl/Ae | 50000.7648 | 4.375 | 30 | 19.1406 | 5
Mass Stability (B1) | M | 34.9995 | kg | 0 | % | 1 | 2 | 0 | 0 | gl/Ae | 50000.7648 | 0 | 30 | 0 | 0
AMH Lubrication (B1a) | M | 0 | kg | 0 | kg | 2 | 1 | 0 | 0 | gl/Ae | 50000.7648 | 0 | 30 | 0 | 0
Piston Mass (B1b) | M | 0.2 | kg | 3e-06 | kg | 2 | 1 | 3e-06 | 1.5e-06 | gl/Ae | 50000.7648 | 0.075001 | 30 | 0.0056252 | 0.085716
Bell Mass (B1a) | M | 0.3 | kg | 5e-06 | kg | 2 | 1 | 5e-06 | 2.5e-06 | gl/Ae | 50000.7648 | 0.125 | 30 | 0.015625 | 0.14286
Local Gravity (B2) | gL | 9.8066 | m/sec2 | 0.0002 | % | 1 | 1 | 1.9613e-05 | 9.8066e-06 | gL-1 x P | 178450.3373 | 1.75 | 30 | 3.0625 | 2
Air Density (B3) | rAIR | 0 | kg/m3 | 0 | kg/m3 | 2 | 1 | 0 | 0 | rM-1 x P | 220.9596 | 0 | 30 | 0 | 0
Ambient Pressure (B3) | PAMB | 101325 | Pa | 200 | Pa | 2 | 1 | 200 | 0 | 1.5E-9 x P | 2.625 | 0 | 30 | 0 | 0
Ambient Temperature (B3) | TAMB | 23 | °C | 1 | °C | 2 | 1 | 1 | 0 | 5.5E-7 x P | 0.9625 | 0 | 30 | 0 | 0
Ambient Humidity (B3) | HAMB | 50 | %RH | 10 | %RH | 2 | 1 | 10 | 0 | 1.2E-8 x P | 0.021 | 0 | 30 | 0 | 0
Mass Load AVG Density (B4) | rM | 7920 | kg/m3 | 40 | kg/m3 | 2 | 1 | 40 | 20 | rAIR/rM2 x P | 0.033479 | 0.66957 | 30 | 0.44833 | 0.76523
Surface Tension (B18) | τ | 0.032 | N/m | 10 | % | 1 | 1 | 0.0032 | 0 | πD/A0 | 253.1238 | 0 | 30 | 0 | 0
Verticality (<0.5 deg.) (B11) | q | 0 | deg. | 0.167 | deg. | 2 | 1 | 0.167 | 0.0835 | 1 - COSΦ x P | 1.8584 | 1.8584 | 30 | 3.4536 | 2.1239
System Terms
Density of Fluid (B6) | rFLUID | 20.1157 | kg/m3 | 0.1 | % | 1 | 1 | 0.020116 | 0.010058 | Aoh/M x P | 0.98067 | 0.0098634 | 30 | 9.7286e-05 | 0.011272
Fluid Head Height (B5) | h | 0.1 | m | 0.01 | m | 2 | 1 | 0.01 | 0.005 | g(rFLUID-rAIR) | 197.2678 | 0.98634 | 30 | 0.97286 | 1.1272
Float Position (B5) | fpi | 0.007 | m | 0.0001 | m | 2 | 1 | 0.0001 | 5e-05 | g(rFLUID-rAIR) | 197.2678 | 0.0098634 | 30 | 9.7286e-05 | 0.011272
Barometer (B8) | Pbaro | 101325 | Pa | 10 | Pa | 2 | 1 | 10 | 0 | 1 | 1 | 0 | 30 | 0 | 0
Reference Pressure, VAC (B7) | PREF | 2 | Pa | 10 | % | 1 | 1 | 0.2 | 0.1 | 1 | 1 | 0.1 | 30 | 0.01 | 0.11429
Diff RPM4 RPT | P | 1750000 | Pa | 0.4 | Pa | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 30 | 0 | 0
This Spreadsheet is provided at no cost and was developed primarily as a learning tool. Fluke Corporation does not guarantee the correctness of the calculations here and does not claim any responsibility for its use. For questions or comments on this spreadsheet please contact Mike Bair at mike.bair@flukecal.com. | [10 empty] | Combined Variance: | 122.9017 | Pa |  | Coverage
[11 empty] | Combined Uncertainty (k=1): | 11.0861 | Pa | k | 2.0086
[11 empty] | Expanded Uncertainty (95%): | 22.2671 | Pa | DOF | 50.4918
[11 empty] | Expanded Uncertainty (95%): | 12.7241 | ppm
Note on B2: Enter maximum pressure in Pa.
Note on I2: Enter the minimum pressure in Pa that you will measure with the Piston Gauge
Note on B7: Enter the maximum head height made with the Piston Gauge.
Note on B8: Enter the p-c effective area reference temperature from the calibration certificate.
Note on B9: Enter the largest temperature correction made for temperature of the piston-cylinder. It is fine to be conservative here. If your lab is 20 to 26 deg C and reference temperature is 23, then this value is 3.
Note on I9: You can use this psi to Pa converter if you are used to using psi. For other units you can just use a simple unit converter tool.
Note on B10: Enter the confidence for the complete uncertainty analysis.
Note on N12: Note that entering 30 for these gives a k of 2. If you are less confident of these uncertainties then reduce these values as appropriate.
Note on A13: The designation at the end of the description is the reference to the PG7000 tech note uncertainty.
Note on E13: Enter the value of the uncertainty for each parameter. It can be a % of the Approx Value column or a value in the unit of measure for that parameter. Enter 1 in the column to the right for % of the parameter value or 2 if it is in the parameter's unit.
Note on C15: Enter the effective area value from the calibration certificate in M^2.
Note on C16: Enter the thermal expansion coefficient for the piston, then cylinder below. If you only know the combination of them you can enter here and use 0 for cylinder.
Note on C18: The reference value plus maximum difference from above is used here. So it is locked.
Note on C19: Enter the pressure distortion coefficient from the certificate here and if there is a 2nd order in the cell below. This must be in 1/MPa. If the report is in another unit use the inverse of the unit conversion. For the 2nd order coefficient use the inverse of the unit conversion squared.
Note on E19: You should only include this uncertainty if you are using a theoretical deformation coefficient as with many PG7000 Piston Gauges.
Note on C21: Always the pressure here so is just what is entered above.
Note on E21: Enter from the PG uncertainty technical Note
Note on E22: You can enter sensitivity in % of value here. Just below you can enter the threshold value. You can find these from the uncertainty technical notes
Note on E24: Drift linked to the Control chart based on the interval assigned and rate of change calculated.
Note on E27: Enter a value for stability here if not included in mass uncertainty above.
Note on E28: Enter uncertainty due to AMH lubrication grease on threads of stem and binary tray from PG9000 technical note.
Note on E29: If piston (or cylinder if floating piece) mass uncertainty is different from main mass enter here.
Note on E30: If bell uncertainty is different from main mass then enter here.
Note on E32: If this is grayed out and not used you have selected absolute by vac mode or diff mode. If you enter zero here when not grayed out then the uncertainty will use the individual uncertainties for the sensors instead of an overall uncertainty in air density.
Note on C36: You should enter worse case average density here. The lowest value is the biggest correction and the highest uncertainty.
Note on E36: It is not intuitive but the uncertainty in mass density gets included when in absolute by vac mode, not gauge.
Note on E38: Convert verticality in minutes to degrees to enter here (minutes/60)
Note on C44: This is limit. Enter the highest vacuum pressure that will be used. If you always wait until vacuum is less than 2 Pa, then enter 2.
Note on C45: What is the contribution of the RPM for diff mode.
Note on E45: What is the uncertainty contribution of the RPM4 due to its repeatability.

## Unccalc2 (hidden)
Limits
Pressure Value | 875000 | Pa | 126.908 | psi | PABS | 875000
Mode | 1 | 1 for ABS, 2 for Gauge, 3 for ABS by ATM
Media :gas | 1 | 1 for gas, 2 for oil, 3 for water |  |  |  | Entry Field
KN (mass to pressure) | 50000.7648 | Pa/kg | 7.252 | psi/kg |  | Not Used
Oil Density | 916 | kg/m3 |  |  |  | Password | Fluke
Maximum Head Height | 0.1 | m | 0.3937 | inches
Reference Temperature | 20 | °C What temperature is the effective area given
P-C Temperature Max Limit | 5 | °C What is the maximum difference in room temperature and reference temperature
Confidence | 95 | %
Influence Quantities |  |  |  | Influence Uncertainties | [5 empty] | Pressure Variant Terms
Influence | Typical | Approx. |  | Uncertainty Value |  | Type Value | Distribution | Uncertainty | Equiv. Std. | Sensitivity | Sensitivity | Std. Unc. | DOF | Variance | Expanded
Description | Symbol | Value | Units | Either % or value in unit 2σ or 95% |  | 1=relative(%), 2=constant | 1=normal, 2=rectangular |  | Unc. (1s) | Coefficient | Value | (Pa) | ----- | (Pa)² | (2σ, PPM)
Effective Area Terms
Effective Area | Ao | 0.00019613 | m2 | 0.0011 | % | 1 | 1 | 2.1574e-09 | 1.0787e-09 | Ao-1 x P | 4461326671.0855 | 4.8125 | 30 | 23.1602 | 11
Piston Thermal Coef. | aP | 4.5e-06 | °C-1 | 2 | % | 1 | 2 | 9e-08 | 5.1962e-08 | (TMAX - TREF) x P | 4375000 | 0.22733 | 30 | 0.05168 | 0.51962
Cylinder Thermal Coef. | aC | 4.5e-06 | °C-1 | 2 | % | 1 | 2 | 9e-08 | 5.1962e-08 | (TMAX - TREF) x P | 4375000 | 0.22733 | 30 | 0.05168 | 0.51962
Piston-Cylinder Temp. | T | 25 | °C | 0.1 | °C | 2 | 1 | 0.1 | 0.05 | (αp+αc) x P | 7.875 | 0.39375 | 30 | 0.15504 | 0.9
1st Order Pressure Coef. | l,b1 | -2.6e-06 | MPa-1 | 10 | % | 1 | 1 | -2.6e-07 | -1.3e-07 | P2 | 765625000000 | -0.099531 | 30 | 0.0099065 | 0.2275
2nd Order Pressure Coef. | b2 | 0 | MPa-2 | 0 | % | 1 | 1 | 0 | 0 | P3 | 6.6992e+17 | 0 | 30 | 0 | 0
Linearity | P | 875000 | Pa | 0.00015 | % | 1 | 1 | 1.3125 | 0.65625 | 1 | 1 | 0.65625 | 30 | 0.43066 | 1.5
Performance | P | 875000 | Pa | 1.5e-05 | % | 1 | 1 | 0.13125 | 0.065625 | 1 | 1 | 0.065625 | 30 | 0.0043066 | 0.15
Sensitivity Limit | P | 875000 | Pa | 0.058 | Pa | 2 | 1 | 0.058 | 0.029 | 1 | 1 | 0.029 | 30 | 0.000841 | 0.066286
Ao Stability (2 yrs) | Ao | 0.00019613 | m2 | -4.6305e-05 | % | 1 | 2 | -9.0818e-11 | -5.2434e-11 | Ao-1 x P | 4461326671.0855 | -0.23392 | 30 | 0.05472 | 0.53468
Force Terms
Mass Load | M | 17.4997 | kg | 0.0005 | % | 1 | 1 | 8.7499e-05 | 4.3749e-05 | gl/Ae | 50000.7648 | 2.1875 | 30 | 4.7852 | 5
Mass Stability | M | 17.4997 | kg | 0 | % | 1 | 2 | 0 | 0 | gl/Ae | 50000.7648 | 0 | 30 | 0 | 0
AMH Lubrication | M | 0 | kg | 0 | kg | 2 | 1 | 0 | 0 | gl/Ae | 50000.7648 | 0 | 30 | 0 | 0
Piston Mass | M | 0.2 | kg | 3e-06 | kg | 2 | 1 | 3e-06 | 1.5e-06 | gl/Ae | 50000.7648 | 0.075001 | 30 | 0.0056252 | 0.17143
Bell Mass | M | 0.3 | kg | 5e-06 | kg | 2 | 1 | 5e-06 | 2.5e-06 | gl/Ae | 50000.7648 | 0.125 | 30 | 0.015625 | 0.28572
Local Gravity | gL | 9.8066 | m/sec2 | 0.0002 | % | 1 | 1 | 1.9613e-05 | 9.8066e-06 | gL-1 x P | 89225.1686 | 0.875 | 30 | 0.76562 | 2
Air Density | rAIR | 0 | kg/m3 | 0 | kg/m3 | 2 | 1 | 0 | 0 | rM-1 x P | 110.4798 | 0 | 30 | 0 | 0
Ambient Pressure | PAMB | 101325 | Pa | 200 | Pa | 2 | 1 | 200 | 0 | 1.5E-9 x P | 1.3125 | 0 | 30 | 0 | 0
Ambient Temperature | TAMB | 23 | °C | 1 | °C | 2 | 1 | 1 | 0 | 5.5E-7 x P | 0.48125 | 0 | 30 | 0 | 0
Ambient Humidity | HAMB | 50 | %RH | 10 | %RH | 2 | 1 | 10 | 0 | 1.2E-8 x P | 0.0105 | 0 | 30 | 0 | 0
Mass Load AVG Density | rM | 7920 | kg/m3 | 40 | kg/m3 | 2 | 1 | 40 | 20 | rAIR/rM2 x P | 0.016739 | 0.33479 | 30 | 0.11208 | 0.76523
Surface Tension | τ | 0.032 | N/m | 10 | % | 1 | 1 | 0.0032 | 0 | πD/A0 | 253.1238 | 0 | 30 | 0 | 0
Verticality (<0.5 deg.) | q | 0 | deg. | 0.167 | deg. | 2 | 1 | 0.167 | 0.0835 | 1 - COSΦ x P | 0.92919 | 0.92919 | 30 | 0.8634 | 2.1239
System Terms
Density of Fluid | rFLUID | 10.0579 | kg/m3 | 0.1 | % | 1 | 1 | 0.010058 | 0.0050289 | Aoh/M x P | 0.98067 | 0.0049317 | 30 | 2.4322e-05 | 0.011272
Fluid Head Height | h | 0.1 | m | 0.01 | m | 2 | 1 | 0.01 | 0.005 | g(rFLUID-rAIR) | 98.6339 | 0.49317 | 30 | 0.24322 | 1.1272
Float Position | fpi | 0.007 | m | 0.0001 | m | 2 | 1 | 0.0001 | 5e-05 | g(rFLUID-rAIR) | 98.6339 | 0.0049317 | 30 | 2.4322e-05 | 0.011272
Barometer | Pbaro | 101325 | Pa | 10 | Pa | 2 | 1 | 10 | 0 | 1 | 1 | 0 | 30 | 0 | 0
Reference Pressure | PREF | 2 | Pa | 10 | % | 1 | 1 | 0.2 | 0.1 | 1 | 1 | 0.1 | 30 | 0.01 | 0.22857
Diff RPM4 RPT | P | 875000 | Pa | 0.4 | Pa | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 30 | 0 | 0
[11 empty] | Combined Variance: | 30.7198 | Pa |  | Coverage
[11 empty] | Combined Uncertainty (k=1): | 5.5425 | Pa | k | 2.0086
[11 empty] | Expanded Uncertainty (95%): | 11.1325 | Pa | DOF | 50.4734
[11 empty] | Expanded Uncertainty (95%): | 12.7229 | ppm

## Unccalc3 (hidden)
Limits
Pressure Value | 25000 | Pa | 3.6259 | psi | PABS | 25000
Mode | 1 | 1 for ABS, 2 for Gauge, 3 for ABS by ATM
Media :gas | 1 | 1 for gas, 2 for oil, 3 for water |  |  |  | Entry Field
KN (mass to pressure) | 50000.7648 | Pa/kg | 7.252 | psi/kg |  | Not Used
Oil Density | 916 | kg/m3 |  |  |  | Password | Fluke
Maximum Head Height | 0.1 | m | 0.3937 | inches
Reference Temperature | 20 | °C What temperature is the effective area given
P-C Temperature Max Limit | 5 | °C What is the maximum difference in room temperature and reference temperature
Confidence | 95 | %
Influence Quantities |  |  |  | Influence Uncertainties | [5 empty] | Pressure Variant Terms
Influence | Typical | Approx. |  | Uncertainty Value |  | Type Value | Distribution | Uncertainty | Equiv. Std. | Sensitivity | Sensitivity | Std. Unc. | DOF | Variance | Expanded
Description | Symbol | Value | Units | Either % or value in unit 2σ or 95% |  | 1=relative(%), 2=constant | 1=normal, 2=rectangular |  | Unc. (1s) | Coefficient | Value | (Pa) | ----- | (Pa)² | (2σ, PPM)
Effective Area Terms
Effective Area | Ao | 0.00019613 | m2 | 0.0011 | % | 1 | 1 | 2.1574e-09 | 1.0787e-09 | Ao-1 x P | 127466476.3167 | 0.1375 | 30 | 0.018906 | 11
Piston Thermal Coef. | aP | 4.5e-06 | °C-1 | 2 | % | 1 | 2 | 9e-08 | 5.1962e-08 | (TMAX - TREF) x P | 125000 | 0.0064952 | 30 | 4.2188e-05 | 0.51962
Cylinder Thermal Coef. | aC | 4.5e-06 | °C-1 | 2 | % | 1 | 2 | 9e-08 | 5.1962e-08 | (TMAX - TREF) x P | 125000 | 0.0064952 | 30 | 4.2188e-05 | 0.51962
Piston-Cylinder Temp. | T | 25 | °C | 0.1 | °C | 2 | 1 | 0.1 | 0.05 | (αp+αc) x P | 0.225 | 0.01125 | 30 | 0.00012656 | 0.9
1st Order Pressure Coef. | l,b1 | -2.6e-06 | MPa-1 | 10 | % | 1 | 1 | -2.6e-07 | -1.3e-07 | P2 | 625000000 | -8.125e-05 | 30 | 6.6016e-09 | 0.0065
2nd Order Pressure Coef. | b2 | 0 | MPa-2 | 0 | % | 1 | 1 | 0 | 0 | P3 | 15625000000000 | 0 | 30 | 0 | 0
Linearity | P | 25000 | Pa | 0.00015 | % | 1 | 1 | 0.0375 | 0.01875 | 1 | 1 | 0.01875 | 30 | 0.00035156 | 1.5
Performance | P | 25000 | Pa | 1.5e-05 | % | 1 | 1 | 0.00375 | 0.001875 | 1 | 1 | 0.001875 | 30 | 3.5156e-06 | 0.15
Sensitivity Limit | P | 25000 | Pa | 0.058 | Pa | 2 | 1 | 0.058 | 0.029 | 1 | 1 | 0.029 | 30 | 0.000841 | 2.32
Ao Stability (2 yrs) | Ao | 0.00019613 | m2 | -4.6305e-05 | % | 1 | 2 | -9.0818e-11 | -5.2434e-11 | Ao-1 x P | 127466476.3167 | -0.0066835 | 30 | 4.467e-05 | 0.53468
Force Terms
Mass Load | M | 0.49999 | kg | 0.0005 | % | 1 | 1 | 2.5e-06 | 1.25e-06 | gl/Ae | 50000.7648 | 0.0625 | 30 | 0.0039062 | 5
Mass Stability | M | 0.49999 | kg | 0 | % | 1 | 2 | 0 | 0 | gl/Ae | 50000.7648 | 0 | 30 | 0 | 0
AMH Lubrication | M | 0 | kg | 0 | kg | 2 | 1 | 0 | 0 | gl/Ae | 50000.7648 | 0 | 30 | 0 | 0
Piston or Cylinder Mass | M | 0.2 | kg | 3e-06 | kg | 2 | 1 | 3e-06 | 1.5e-06 | gl/Ae | 50000.7648 | 0.075001 | 30 | 0.0056252 | 6.0001
Bell Mass | M | 0.3 | kg | 5e-06 | kg | 2 | 1 | 5e-06 | 2.5e-06 | gl/Ae | 50000.7648 | 0.125 | 30 | 0.015625 | 10.0002
Local Gravity | gL | 9.8066 | m/sec2 | 0.0002 | % | 1 | 1 | 1.9613e-05 | 9.8066e-06 | gL-1 x P | 2549.2905 | 0.025 | 30 | 0.000625 | 2
Air Density | rAIR | 0 | kg/m3 | 0 | kg/m3 | 2 | 1 | 0 | 0 | rM-1 x P | 3.1566 | 0 | 30 | 0 | 0
Ambient Pressure | PAMB | 101325 | Pa | 200 | Pa | 2 | 1 | 200 | 0 | 1.5E-9 x P | 0.0375 | 0 | 30 | 0 | 0
Ambient Temperature | TAMB | 23 | °C | 1 | °C | 2 | 1 | 1 | 0 | 5.5E-7 x P | 0.01375 | 0 | 30 | 0 | 0
Ambient Humidity | HAMB | 50 | %RH | 10 | %RH | 2 | 1 | 10 | 0 | 1.2E-8 x P | 0.0003 | 0 | 30 | 0 | 0
Mass Load AVG Density | rM | 7920 | kg/m3 | 40 | kg/m3 | 2 | 1 | 40 | 20 | rAIR/rM2 x P | 0.00047827 | 0.0095654 | 30 | 9.1496e-05 | 0.76523
Surface Tension | τ | 0.032 | N/m | 10 | % | 1 | 1 | 0.0032 | 0 | πD/A0 | 253.1238 | 0 | 30 | 0 | 0
Verticality (<0.5 deg.) | q | 0 | deg. | 0.167 | deg. | 2 | 1 | 0.167 | 0.0835 | 1 - COSΦ x P | 0.026548 | 0.026548 | 30 | 0.00070482 | 2.1239
System Terms
Density of Fluid | rFLUID | 0.28737 | kg/m3 | 0.1 | % | 1 | 1 | 0.00028737 | 0.00014368 | Aoh/M x P | 0.98067 | 0.00014091 | 30 | 1.9854e-08 | 0.011272
Fluid Head Height | h | 0.1 | m | 0.01 | m | 2 | 1 | 0.01 | 0.005 | g(rFLUID-rAIR) | 2.8181 | 0.014091 | 30 | 0.00019854 | 1.1272
Float Position | fpi | 0.007 | m | 0.0001 | m | 2 | 1 | 0.0001 | 5e-05 | g(rFLUID-rAIR) | 2.8181 | 0.00014091 | 30 | 1.9854e-08 | 0.011272
Barometer | Pbaro | 101325 | Pa | 10 | Pa | 2 | 1 | 10 | 0 | 1 | 1 | 0 | 30 | 0 | 0
Reference Pressure | PREF | 2 | Pa | 10 | % | 1 | 1 | 0.2 | 0.1 | 1 | 1 | 0.1 | 30 | 0.01 | 8
Diff RPM4 RPT | P | 25000 | Pa | 0.4 | Pa | 2 | 1 | 0 | 0 | 1 | 1 | 0 | 30 | 0 | 0
[11 empty] | Combined Variance: | 0.057135 | Pa |  | Coverage
[11 empty] | Combined Uncertainty (k=1): | 0.23903 | Pa | k | 1.9784
[11 empty] | Expanded Uncertainty (95%): | 0.47289 | Pa | DOF | 130.5245
[11 empty] | Expanded Uncertainty (95%): | 18.9156 | ppm

## Historical_Ae
Due Date | Aging Rate/ Ae/day | Y intercept | Days Until Due Date | Number of days from last Cal | Number of Cals | Sum Square of Calibration Time | Sum of calibration time squared | Computation of Sxx | Mean calibration time | Standard Error of the Fit | Calibration Interval (mos)
2017-05-12 | -2.0729e-08 | 49.0199 | -489.5786 | 1584.5786 | 9 | 81167002 | 467164996 | 29259780.2222 | 2401.5556 | 0.00016064 | 36
|  | Confidence based on todays date and next cal date | 0.00041365 | 0.00043101 | m²
Mass Calibration History |  |  | 8.4385 | 8.7926 | PPM
Date | Effective Area | Ae UNC | Ae UNC | Days | Pred. Value | Upper bound | Lwr bound | Predicted - Actual |  | Forecast Ae Value Next Cal
| Value | 95% (k=2) | 95% (k=2) | [6 empty] | 49.019
| [m²] | [PPM] | [m²] |  | [kg] | [kg] | [kg] | [PPM] |  | Forecast Uncertainty
1999-11-02 | 49.0192 | 16 | 0.00078431 | 0 | 49.0192 | 49.0195 | 49.0188 | 0.0205 |  | 8.4385
2000-11-09 | 49.0194 | 16 | 0.00078431 | 373 | 49.0192 | 49.0195 | 49.0188 | -5.6657 |  | Difference in Forecast and last
2001-11-12 | 49.0189 | 16 | 0.0007843 | 741 | 49.0192 | 49.0195 | 49.0188 | 4.8276 |  | -2.36 | PPM
2003-11-12 | 49.0191 | 10 | 0.00049019 | 1471 | 49.0191 | 49.0195 | 49.0188 | 0.17365 |  | 17
2005-11-14 | 49.0192 | 10 | 0.00049019 | 2204 | 49.0191 | 49.0195 | 49.0188 | -1.1155 |  | Stability rate based on interval in %
2008-05-05 | 49.0189 | 8 | 0.00039215 | 3107 | 49.0191 | 49.0194 | 49.0188 | 3.7659 |  | -4.63e-05%
2010-05-07 | 49.0191 | 8 | 0.00039215 | 3839 | 49.0191 | 49.0194 | 49.0187 | 0.53907
2012-05-10 | 49.0191 | 8 | 0.00039215 | 4573 | 49.0191 | 49.0194 | 49.0187 | -0.64853
2014-05-13 | 49.0192 | 8 | 0.00039215 | 5306 | 49.0191 | 49.0194 | 49.0187 | -1.8969
Formulas:
Due Date = MAX(A9:A26)+[Calibration Interval (mos)]/12*365

## Historical_mass
Due Date | Aging Rate/ grams/day | Y intercept | Days Until Due Date | Number of days from last Cal | Number of Cals | Sum Square of Calibration Time | Sum of calibration time squared | Computation of Sxx | Mean calibration time | Standard Error of the Fit | Calibration Interval (mos)
2017-05-12 | -1.468e-06 | 200.044 | -489.5786 | 1584.5786 | 9 | 41927268 | 251412736 | 13992519.5556 | 1761.7778 | 0.00057898 | 36
|  | Confidence based on todays date and next cal date | 0.0015658 | 0.0016649 | kg
Mass Calibration History |  |  | 7.8273 | 8.3228 | PPM
Date | Mass |  | Mass UNC | Days | Pred. Value | Upper bound | Lwr bound | Predicted - Actual |  | Forecast Mass Value Next Cal
| Value |  | 95% (k=2) | [6 empty] | 199.9811 | grams
| [g] |  | [g] |  | [g] | [g] | [g] | [g] |  | Forecast Uncertainty
2003-11-12 | 199.989 |  | 0.003 | 0 | 199.9884 | 199.9897 | 199.987 | -0.00064708 |  | 0.0015658 | grams
2004-11-05 | 199.9876 |  | 0.003 | 359 | 199.9878 | 199.9891 | 199.9865 | 0.00022592 |  | Difference in Forecast and last
2005-11-14 | 199.9874 |  | 0.003 | 733 | 199.9873 | 199.9885 | 199.986 | -0.00012311 |  | -0.0023842 | grams
2006-11-09 | 199.987 |  | 0.003 | 1093 | 199.9867 | 199.988 | 199.9855 | -0.00025158 |  | 17
2008-05-05 | 199.9853 |  | 0.003 | 1636 | 199.986 | 199.9872 | 199.9847 | 0.00065131 |  | Stability rate based on interval in grams
2010-05-07 | 199.984 |  | 0.003 | 2368 | 199.9849 | 199.9861 | 199.9836 | 0.00087674 |  | -0.0016074
2011-05-04 | 199.9843 |  | 0.003 | 2730 | 199.9843 | 199.9856 | 199.9831 | 4.5334e-05
2012-05-10 | 199.9838 |  | 0.003 | 3102 | 199.9838 | 199.9851 | 199.9825 | -7.5541e-07
2014-05-13 | 199.9835 |  | 0.003 | 3835 | 199.9827 | 199.9841 | 199.9813 | -0.00077679
Formulas:
Due Date = MAX(A9:A26)+[Calibration Interval (mos)]/12*365

## Revision
REVISION LOG: Piston Gauge Uncertainty Analysis
Revision Number | Date | Comments | Author | Function
1 | 20151006 | Initial creation of spreadsheet. | Michael Bair | Metrologist
2 | 20151118 | Fixed the format of the Control charts | Michael Bair | Metrologist
3 | 20170302 | Added Diff mode and Description Field | Michael Bair | Metrologist
4 | 20170303 | Linked stability of piston Ae and mass to control charts | Michael Bair | Metrologist