COMPASS for Pressure Calibration Software

Global macro to specify Reference uncertainty by range in kPa

Macro code to ignore the uncertainty values that you have entered for your references (support devices and/or piston gauges) in COMPASS for Pressure software. This macro allows you to specify your reference uncertainty (in % Reading and/or % Span) according to ranges of pressure. This might correlate to having different standards for different ranges of pressures. This can be used with “Measurement Uncertainty in COMPASS for Pressure - Test Macro” and would be inserted after the degrees of freedom check, before the ExpUnc calculation.

Copy and Paste all of the following text into a new Global Macro.

'Alternate Reference Uncertainty Calculation/Determination - by Pressure Range

'converting RefP to kPa for range check

Function UncRefByRange(RefF, RefFunit, RefFunitTxt)

i=1 'use For -> Next loop for multiple DUTs

dUnit = cCOMPASS.cConfig.DUTPrs(i).RangeMain.UnitFinal

cDebug.LogStatus "DUT Units = " & dUnit

RefP_kPa = cCOMPASS.UnitConversion(CDbl(RefP), 3, CInt(dUnit), 0) '3 is integer ID for kPa

cDebug.LogStatus "RefP_kPa = " & RefP_kPa

'range specific coefs for calculation of reference uncertainty

Delta = 0

If RefP_kPa >= 7000 Then		'kPa, A20Ms

PctRdg = 0.01/100

ElseIf RefP_kPa >= 3500 Then	'A7Ms

PctRdg = 0.01/100

ElseIf RefP_kPa >= 700 Then		'A3.5Ms

PctRdg = 0.01/100

ElseIf RefP_kPa >= 200 Then 	'A700K

PctRdg = 0.01/100

ElseIf RefP_kPa >=60 Then 		'A200K

PctRdg = 0.01/100

Else

PctRdg = 0.0			'A200K

Delta = 21				'Threshold Unc.

End If

cDebug.LogStatus "PctRdg = " & PctRdg

cDebug.LogStatus "Delta = " & Delta & " " & "kPa"

'converting Delta to test unit

Delta = cCOMPASS.UnitConversion(CDbl(Delta), CInt(dUnit),1, 0)

cDebug.LogStatus "Delta = " & Delta & " " & dUnit

URef = (RefP*PctRdg) + Delta

cDebug.LogStatus "Reference Tolerance (alt) = " & Uref

End Function