COMPASS for Pressure with 2271A, 6270A, 8270A or 8370A Pressure Controller

Reset Safety Limits at End of Test, and/or

Set Hold Limit to Pressure Unit Mode (not %Range)

Applies to 2271A, 6270A, 8270A and 8370A pressure controllers

If you want to reset the Safety Limits and Ready Tolerance settings at the end of a COMPASS for Pressure test , then run a “Post Test Macro” named “SafetyLimitsRST” at the end of the test. Add a post test macro in the Data tab in the “Test Editor” window of an Advanced Test. Create this macro if necessary in the COMPASS Macro Editor and name it “SafetyLimitsRST”.

Here’s a version that simply uses the *RST command

Function SafetyLimitsRST(Target, ChnlSt, SetMode, ParamID, cDevice)

rtv = cCOMPASS.cConfig.ActiveRefP.IoSendComman(Cstr(*RST),True)

End Function

Here’s an alternate version

Function SafetyLimitsRST(Target, ChnlSt, SetMode, ParamID, cDevice)

Cmd1 = "UNIT:PRES PSI"

Cmd2 = "CALCulate:LIMit:LOWer -300"

Cmd3 = "CALCulate:LIMit:UPPer 3100"

Cmd4 = "SOURce:PRESsure:TOLerance:RANGe On"

Cmd5 = "SOURce:PRESsure:TOLerance 0.005"

For i = 1 To 5

num = "Cmd" & i

cmd = num

rtv = cCOMPASS.cConfig.ActiveRefP.IoSendComman(Cstr(cmd),True)

Next

End Function

Another alternate version is on page 2

'**************************************************

'Test Macros do not have a return value.

'Manipulate the test or device collection as desired.

'iT  The current temperature point in the test

'iL  The current line pressure point in the test.

'iC  The current pressure cycle in the test.

'IP  The current pressure point in the test.

'cTest   The test class .

'cConfig  Configuration class that holds all active devices.

'

'**************************************************

Function PostTest6270A(iT, iL, iC, iP, cTest, cConfig)

'Reset 6270A to default settings

'Will also work without changes for 2271A

'Works with 8270A and 8370A but values need to be changed

cmd = "UNIT:PRES PSI"

cCOMPASS.StatusDisplay "Resetting 6270A default settings"

reply = cCOMPASS.cConfig.SetPrs(1).ioSendCommand(CStr(cmd),True,0)

cDebug.LogStatus "Sent command: " & cmd

cmd = "CALC:LIMIT:UPP 3300"

cCOMPASS.StatusDisplay "Resetting 6270A default settings"

reply = cCOMPASS.cConfig.SetPrs(1).ioSendCommand (CStr(cmd),True,0)

cDebug.LogStatus "Sent command: " & cmd

cmd = "CALC:LIMIT:LOW -300"

cCOMPASS.StatusDisplay "Resetting 6270A default settings"

reply = cCOMPASS.cConfig.SetPrs(1).ioSendCommand (CStr(cmd),True,0)

cDebug.LogStatus "Sent command: " & cmd

'change 6270A to most common unit that is used and % of range for hold/tolerance limit

cmd = "UNIT:PRES PSI"

cCOMPASS.StatusDisplay "Resetting 6270A default settings"

reply = cCOMPASS.cConfig.SetPrs(1).ioSendCommand (CStr(cmd),True,0)

cDebug.LogStatus "Sent command: " & cmd

cmd = "TOL:RANGE ON"

cCOMPASS.StatusDisplay "Resetting 6270A default settings"

reply = cCOMPASS.cConfig.SetPrs(1).ioSendCommand (CStr(cmd),True,0)

cDebug.LogStatus "Sent command: " & cmd

cmd = "TOL 0.005"

cCOMPASS.StatusDisplay "Resetting 6270A default settings"

reply = cCOMPASS.cConfig.SetPrs(1).ioSendCommand (CStr(cmd),True,0)

cDebug.LogStatus "Sent command: " & cmd

TimeDelay 1

cCOMPASS.StatusDisplay "Resetting 6270A default settings"

msgbox "Remember to close COMPASS before connecting to another pressure module."

cCOMPASS.StatusDisplay "Press OK to close message box"

TimeDelay 1

End Function

If you want to set the Ready Tolerance settings to Pressure Unit mode (not % Range mode) at the beginning of a COMPASS for Pressure test , then run a “Pre Test Macro” named “PreTest6270A” at the beginning of the test. Create this macro if necessary in the COMPASS Macro Editor and name it “PreTest6270A.

'**************************************************

'Test Macros do not have a return value.

'Manipulate the test or device collection as desired.

'iT  The current temperature point in the test

'iL  The current line pressure point in the test.

'iC  The current pressure cycle in the test.

'IP  The current pressure point in the test.

'cTest   The test class .

'cConfig  Configuration class that holds all active devices.

'

'**************************************************

Function PreTest6270A(iT, iL, iC, iP, cTest, cConfig)

'Will also work without changes for 2271A, 8270A, 8370A

'change 6270A to pressure unit for hold/tolerance limit

cmd = "TOL:RANGE OFF"

cCOMPASS.StatusDisplay "Set 6270A hold limit to pressure units, not % of range"

reply = cCOMPASS.cConfig.SetPrs(1).ioSendCommand (CStr(cmd),True,0)

cDebug.LogStatus "Sent command: " & cmd

TimeDelay 1

End Function