|  | Calibration Report mA test with slope and offset correction Model: PMX459-5.0KGI Serial Number: 425595 | Nov 15 2015 |
| --- | --- | --- |

Report Information

Report Compiled:Nov 15 2015

Date of test:20140520

Time of test:3:22:46 PM

Data file:C:\20140520_002.dat

Test file:mA test with slope and offset correction

| DUT Information |  |  | Reference Information |  |
| --- | --- | --- | --- | --- |
| Manufacturer | Omegadyne |  | Manufacturer | COMPASS |
| Model | PMX459-5.0KGI |  | Model | Manual Device |
| Serial Number | 425595 |  | Serial Number | N/A |
| Pressure Range | 0.000 to 5000.000 psi |  | Pressure Range | 0.000 to 100.000 psi |
| Nominal Uncertainty | 0.08 %Span |  | Nominal Uncertainty |  |

Test Data

| Set Pt | Reference Pressure | DUT Pressure | DUT Output | Abs. Error | "% Span" Error | DUT Tolerance | Status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| psi | psi | psi | mA | psi | % | psi |  |
| 0.0000 | 0.0000 | 6.2 | 4.020 | 6.2 | 0.1240 | 4.0 | T |
| 2500.0000 | 2500.0000 | 2511.6 | 12.037 | 11.6 | 0.2320 | 4.0 | T |
| 5000.0000 | 5000.0000 | 5015.9 | 20.051 | 15.9 | 0.3180 | 4.0 | T |
| 2500.0000 | 2500.0000 | 2512.5 | 12.040 | 12.5 | 0.2500 | 4.0 | T |
| 0.0000 | 0.0000 | 6.9 | 4.022 | 6.9 | 0.1380 | 4.0 | T |

Calibration Equation: Pressure = Slope * DUT Raw Output + Offset

Slope: 311.9012 psi per mA

Offset: 4.021714 mA