COMPASS for Pressure (or Flow)

Alert if ambient conditions (pressure, temperature, relative humidity) are outside limits

Create a test macro to set ambient limits and specify what to do if conditions are outside the limits

'**************************************************

'Test Macros do not have a return value.

'Manipulate the test or device collection as desired.

'iT  The current temperature point in the test

'iL  The current line pressure point in the test.

'iC  The current pressure cycle in the test.

'IP  The current pressure point in the test.

'cTest   The test class .

'cConfig  Configuration class that holds all active devices.

'

'**************************************************

Function AmbLimitMonitor(iT, iL, iC, iP, cTest, cConfig)

'enter e-mail address if using the e-mail notification feature

'The method used relies on MS Outlook(R) being open when running the test

'comment out the email lines if not using this feature

emailto = "noreply@flukecal.com"

Delay = 900 '(sec), time to wait between email notifications

'enter limits and units as desired

PrsAmbLL = 85

PrsAmbHL = 115

PrsAmbLtxt = "kPa"	'Set this to mbar, kPa, psi or  hPa

PrsAmbLunit = 3		'Set to corresponding integer: Pa=1, mbar=2, kPa=3

'mmHG0C=6, Torr=7, psi=9, inHg32F=26, hPa=38

'	cDebug.LogStatus "Amb Pressure Low Limit: " & PrsAmbLL & " " & PrsAmbLtxt & " (" & PrsAmbLunit & ")"

TempAmbLL = 15

TempAmbHL = 25

TempAmbLtxt = "C"	'Set to C or F

TempAmbLunit = 0	'Set to corresponding integer; C=0, F=1, K=2, R=3

RH_AmbLL = 20 '%

RH_AmbHL =  70 '%

RH_AmbLtxt = "%"

PrsAmb = cCompass.cConfig.AmbPrs.FinalOutput

PrsAmbU = cCompass.cConfig.AmbPrs.RangeMain.UnitFinal

PrsAmbUtxt = cCompass.cConfig.AmbPrs.RangeMain.UnitFinalText

'	cDebug.LogStatus "AmbPrs = " & PrsAmb & " " & PrsAmbUtxt & " (" & PrsAmbU & ")"

PrsLL = cCOMPASS.UnitConversion (CDbl(PrsAmbLL),CInt(PrsAmbU),CInt(PrsAmbLunit),0)

'								(input,newUnit,curUnit,outputType: 0-pressure, 1-temperature, 10-Humidity

'	cDebug.LogStatus "AmbPrsLL = " & PrsLL & " " & PrsAmbUtxt & " (" & PrsAmbU & ")"

PrsHL = cCOMPASS.UnitConversion (CDbl(PrsAmbHL),CInt(PrsAmbU),CInt(PrsAmbLunit),0)

'	cDebug.LogStatus "AmbPrsHL = " & PrsHL & " " & PrsAmbUtxt & " (" & PrsAmbU & ")"

If PrsAmb < PrsLL  Then

msg = "Amb Pressure is too low: " & PrsAmb & " " & PrsAmbUtxt

ambAlert msg,emailto,delay

ElseIf PrsAmb > PrsHL Then

msg = "Amb Pressure is too high: " & PrsAmb & " " & PrsAmbUtxt

ambAlert msg,emailto,delay

End If

TempAmb = cCompass.cConfig.AmbTemp.FinalOutput

TempAmbU = cCompass.cConfig.AmbTemp.RangeMain.UnitFinal

TempAmbUtxt = cCompass.cConfig.AmbTemp.RangeMain.UnitFinalText

'	cDebug.LogStatus "AmbTemp " & TempAmb & " " & TempAmbUtxt & " (" & TempAmbU & ")"

TempLL = cCOMPASS.UnitConversion (CDbl(TempAmbLL),CInt(TempAmbU),CInt(TempAmbLunit),1)

'	cDebug.LogStatus "AmbTempLL = " & TempLL & " " & TempAmbUtxt & " (" & TempAmbU & ")"

TempHL = cCOMPASS.UnitConversion (CDbl(TempAmbHL),CInt(TempAmbU),CInt(TempAmbLunit),1)

'	cDebug.LogStatus "AmbTempHL = " & TempHL & " " & TempAmbUtxt & " (" & TempAmbU & ")"

If TempAmb < TempLL  Then

msg = "Amb Temperature is too low: " & TempAmb & " " & TempAmbUtxt

ambAlert msg,emailto,delay

ElseIf TempAmb > TempHL Then

msg = "Amb Temperature is too high: " & TempAmb & " " & TempAmbUtxt

ambAlert msg,emailto,delay

End If

RH_Amb = cCompass.cConfig.AmbRH.FinalOutput

RH_AmbUtxt = RH_AmbLtxt

'	cDebug.LogStatus "AmbRH = " & RH_Amb & RH_AmbUtxt

'	cDebug.LogStatus "AmbRH_HL = " & RH_AmbHL & RH_AmbLtxt

'	cDebug.LogStatus "AmbRH_LL = " & RH_AmbLL & RH_AmbLtxt

If RH_Amb < CDbl(RH_AmbLL)Then

msg = "Amb RH is too low: " & RH_Amb & " " & RH_AmbUtxt

ambAlert msg,emailto,delay

ElseIf RH_Amb > CDbl(RH_AmbHL) Then

msg = "Amb RH is too high: " & RH_Amb & " " & RH_AmbUtxt

ambAlert msg,emailto,delay

End If

End Function

Function ambAlert (msgAlert,toEmail,delay)

msgAlert = msgAlert  & ". Waiting " & formatnumber (delay/60,1)& " minutes until checking again. "

'	cCOMPASS.StatusDisplay msgAlert

cDebug.LogStatus msgAlert

msgAlert = msgAlert & "Current Date/time is " & Date & ", " & time

msgbox msgAlert, vbCritical, "Alert"

subject = msgAlert

SendEmail toEmail, subject, msgAlert

cDebug.LogStatus "Sent email to: " & toEmail

TimeDelayWstatus (delay)

End Function

The above macro calls the Global macro TimeDelayWstatus. TimeDelayWStatus should be added as a Global macro

'****************************************

'Delay for the amount of time in seconds

'specified by tD. Use the COMPASS object

'to allow for asynchronous operation

'****************************************

Sub TimeDelayWstatus(tD)

tStart = Timer'Timer function returns the number of seconds since 12:00 AM

Do

cCOMPASS.TimeDelay 1

tDiff = Time_Difference(tStart)

remainT = CInt(tD - tDiff)

cCOMPASS.StatusDisplay "Delay remaining..." & remainT

If cCOMPASS.SystemAbort Then Exit Sub

If remainT <= 0 Then Exit Do

Loop Until False

End Sub

And this macro calls the Global macro TimeDelay. It is already in COMPASS as a Global macro but the text is here if you need to add it.

'****************************************

'Delay for the amount of time in seconds

'specified by tD. Use the COMPASS object

'to allow for asynchronous operation

'****************************************

Sub TimeDelay(tD)

cCOMPASS.TimeDelay tD

End Sub

Same with Global macro Time_Difference

'///////////////////////////////////////////////////////////////

'Return the time difference from the current

'time and the time value passed T:

'

'In the event the diff is <0 then add the

'number of seconds in a day because midnight

'has passed since the start time.

'

'NOTE

'This function as an effective time resoluteion

'of 10 ms.

'

'///////////////////////////////////////////////////////////////

Public Function Time_Difference(t)

Dim T_Diff

T_Diff = Timer - t

If T_Diff < -0.1 Then

T_Diff = T_Diff + 86400

ElseIf T_Diff < 0 Then

T_Diff = 0

End If

Time_Difference = T_Diff

End Function