COMPASS for Pressure

Test Macro to calculate real time per point hysteresis and average error

Make sure UserDefined1, UserDefined2 and UserDefined3 are saved to the data file in [Tools], <Options>, <Data in File>

Copy and Paste all of the following text to a Test Macro in the COMPASS Macro Editor

'**************************************************

'For help with the programming syntax,

'search for keyword 'vbscript' on the internet

'**************************************************

'Test Macros do not have a return value, but can save data as specified

'Manipulate the test or device collection as desired

'iT  The current temperature point in the test

'iL  The current line pressure point in the test

'iC  The current pressure cycle in the test

'IP  The current pressure point in the test

'cTest   The test class

'cConfig  Configuration class that holds all active devices

'

'This test macro calculates hysteresis and average error for each descending setpoint

'in the test with an odd number of setpoings (it doesn't work for even number of setpoints)

'The descending setpoints must be the same as the ascending setpoints and the highest point

'cannot be repeated. The macro also calculates DUT error as “Ref – DUT” and saves that to UserDefined3

'

'Make sure UserDefined1, UserDefined2 and UserDefined3 are saved to the data file in

'[Tools], <Options>, <Data In File>, and ensure that they are not being used in other macros or

'by regular COMPASS features (e.g. when using actual flow units COMPASS for Flow saves

'pressure and temperature to Userdefined1 and UserDefined2

'Can only use this macro in Advanced tests, and specify this as the "Test Event Macro"

'on the Data tab, and "Timed Macro Interval(ms)" needs to be 0 (zero)

'

'It is recommended to manually verify the calculated values before relying solely on this macro

'Contact Pressure/Flow Technical Support at pressuresupport@flukecal.com with questions

'**************************************************

Function HysPerPoint(iT, iL, iC, iP, cTest, cConfig)

'Renaming column headers and set resolution of Hysteresis and AvgErr

If cCOMPASS.CurrentTestStep = 1100 Then 'Do this at start of a new pressure/flow cycle

cDebug.LogStatus "*"	'blank line in macro spy window

cDebug.LogStatus "New Test"

For i = 1 To cCOMPASS.DataCollection.Count

With cCOMPASS.DataCollection(i).DataInfoCollection("ID410001")

.Nomenclature = "Hysteresis"	'usually UserDefined1

'.Unit = "% rdg"

'.DataFormat = "0.000E-0"	'Alternate fixed way to specify resolution

'.DataFormat = "0.000"	'Alternate fixed way to specify resolution

DUTrez = CInt(cConfig.DUTPrs(1).RangeMain.ResFinal)

'				cDebug.LogStatus "DUT Resolution (DUTrez) = " & DUTrez

HysRez = CInt(DUTrez)

'.DataFormat = DUTrez

'       		    cDebug.LogStatus "Hysteresis Resolution (HysRez) = " & HysRez

RefRez = CInt(cConfig.RefPrs(1).RangeMain.ResFinal)

'.DataFormat = DUTrez

'       		    cDebug.LogStatus "Reference Resolution (RefRez) = " & RefRez

End With

With cCOMPASS.DataCollection(i).DataInfoCollection("ID410002")

.Nomenclature = "AvgErr"	'usually UserDefined2

AvgErrRez = DUTrez

'.DataFormat = DUTrez

'     		cDebug.LogStatus "AvgErr Resolution (AvgErrRez) = " & AvgErrRez

End With

With cCOMPASS.DataCollection(i).DataInfoCollection("ID410003")

.Nomenclature = "Ref - DUT"	'usually UserDefined3

Ref_DUTrez = DUTrez

'.DataFormat = DUTrez

'      		cDebug.LogStatus "Ref - DUT Resolution (Ref_DUTrez) = " & Ref_DUTrez

End With

cCOMPASS.DataCollection(i).ReWriteFileHeader

Next

End If

If cCompass.CurrentTestStep = 325 Then 	'Averaging is completed but not yet written to data file

For i = 1 To cConfig.DUTPrs.Count

DUTmodel = cConfig.DUTPrs(CInt(i)).RangeMain.GetParent.Model

DUTsn = cConfig.DUTPrs(CInt(i)).RangeMain.GetParent.SN

'		cDebug.LogStatus "DUT Model " & DUTmodel & " SN " & DUTsn

dUnit = cCOMPASS.cConfig.DUTPrs(i).RangeMain.UnitFinal

'		cDebug.LogStatus "DUT Units = " & dUnit

dUnitText = cCOMPASS.cConfig.DUTPrs(i).RangeMain.UnitFinalText

'		cDebug.LogStatus "DUT Unit Text = " & dUnitText

' Change DUT error value from (DUT-Ref) to (Ref-DUT)

'		cDebug.LogStatus "DUT - Ref (from data collection) = " & cCOMPASS.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).DUTDiff & " " & dUnitText

DUTErr = 0 - cCOMPASS.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).DUTDiff

'		cDebug.LogStatus "DUTErr (Ref - DUT) = " & DUTErr & " " & dUnitText

'Get inputs for Hysteresis and AvgErr calculation

DutPrs = cCOMPASS.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).DUTPressure

'		cDebug.LogStatus "DUT Pressure = " & DutPrs

DutRes = cConfig.DUTPrs(i).RangeMain.ResFinal

'		cDebug.LogStatus "DUT Res Rtv (DutRes)= " & DutRes

Res = 0.1^DutRes

'		cDebug.LogStatus "DUT Res (Res) = " & Res

n = cCOMPASS.cCurTest.PointP.count

'    	cDebug.LogStatus "Number of points in test (n) = " & n

n_2 = CDbl(n/2)

n_2 = CStr(n_2)

'   	cDebug.LogStatus "n/2 = " & n_2 & " (CDbl format)"

h1 = CInt(n_2 - 0.49999999999)

'		cDebug.LogStatus "h1 (int of n/2) = " & h1

h = CInt(h1 + 2)

'		cDebug.LogStatus "h, Start hysteresis calc at point (h) =  " & h

p = CInt(cCOMPASS.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).Point)

'		cDebug.LogStatus "*"	'blank line in macro spy window

'		cDebug.LogStatus "Point (p) = " & p

A = CInt(n - p + 1)

'		cDebug.LogStatus "AscPt (n-p+1) = " & A

If p > CInt(h-1) Then

'		cDebug.LogStatus "Do While p > (h-1)"

'				cDebug.LogStatus "DUT Error = " & DUTErr & " " & dUnitText

ErrAscPt = CDbl(-(cCOMPASS.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(A)).DUTDiff))

'				cDebug.LogStatus "Err Asc Pt (" & A & ") = " & ErrAscPt & " " & dUnitText

hys = CDbl(ErrAscPt - DUTErr)

'				cDebug.LogStatus "hys = " & hys

'				cDebug.LogStatus "hys after format with HysRez(" & HysRez & ") = " & formatNumber(hys,HysRez)

AvgErr = CDbl(DUTErr + ErrAscPt)/2

'				cDebug.LogStatus "Avg Err = " & AvgErr

'				cDebug.LogStatus "Avg Err after format with AvgErrRez(" & AvgErrRez & ") = " & formatNumber(hys,AvgErrRez)

cCompass.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).UserDefined1 = Hys	'Hysteresis to UserDefined1

'				cDebug.LogStatus "Collection " & i & ", point " & iP & ", UserDefined1 (Hys) = " & cCompass.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).UserDefined1

cCompass.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).UserDefined2 = AvgErr  'AvgErr to UserDefined2

'				cDebug.LogStatus "Collection " & i & ", point " & iP & ", UserDefined2 (Avg Err) = " & cCompass.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).UserDefined2

Else If (p < h)Then		'Case xxx 'calculate hysteresis only on points past the 1/2 way point of the test

cDebug.LogStatus "Do While p < h"

'			hys = Nothing

'			hys = Null

hys = ""

AvgErr = ""

End If

End If

'		cDebug.LogStatus "* DUT Model " & DUTmodel & ", SN " & DUTsn

'		cDebug.LogStatus "* Ref Pressure = " & cCOMPASS.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).RefPressure

'		cDebug.LogStatus "* DUT Pressure = " & formatNumber(DutPrs,DUTrez) & " " & dUnitText

'		cDebug.LogStatus "* Asc Pt (A) " & A & " Pressure = " & DutPrs & " " & dUnitText

'		cDebug.LogStatus "* Ref-DUT = " & DUTErr & " " & dUnitText

'		cDebug.LogStatus "* Hys (Ref, Up - Down) = " & hys & " " & dUnitText

'		cDebug.LogStatus "* Avg Err (Up, Down) = " & AvgErr & " " & dUnitText

cCompass.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).UserDefined3 = DUTErr  'Rewrite with "Ref - DUT"

'		cDebug.LogStatus "Collection " & i & ", point " & iP & ", UserDefined3 (Ref-DUT) = " & cCompass.DataCollection(i).DataPointRef (1,1,CInt(iC),CInt(iP)).UserDefined3

Next

End If

End Function

'Author: Kyle Clark

'Rev 1, 18-August-2017