COMPASS for Pressure

Test Macro to calculate real time expanded uncertainty with guardband

Make sure UserDefined8, UserDefined9, UserDefined10, StringData8 and StringData9 are saved to the data file in [Tools], <Options>, <Data in File>

Copy and Paste all of the following text to a Test Macro in the COMPASS Macro Editor

'**************************************************

'For help with the programming syntax,

'search for keyword 'vbscript' on the internet

'**************************************************

'Test Macros do not have a return value, but can save data as specified

'Manipulate the test or device collection as desired

'iT  The current temperature point in the test

'iL  The current line pressure point in the test

'iC  The current pressure cycle in the test

'IP  The current pressure point in the test

'cTest   The test class

'cConfig  Configuration class that holds all active devices

'

'This test macro calculates uncertainty and a guardbanded tolerance for each setpoint in the test

'Calculation is done per Fluke Calibration Application Note -

'"Implementing ISO 17025 Measurement Uncertainty Requirements In Software"

'dated March 2012, document 1282496C

'Seach www.flukecal.com for "1282496C" to find this App Note

'

'Revision history at bottom of document

'Make sure UserDefined8, UserDefined9, & StringData9 are saved to the data file in

'[Tools], <Options>, <Data In File>, and ensure that they are not being used in other macros or

'by regular COMPASS features (e.g. when using actual flow units COMPASS for Flow saves

'pressure and temperature to Userdefined1 and UserDefined2

'Can only use this macro in Advanced tests, and specify this as the "Test Event Macro"

'on the Data tab, and "Timed Macro Interval(ms)" needs to be 0 (zero)

'

'It is recommended to manually verify the calculated values before relying solely on this macro

'Contact Pressure/Flow Technical Support at pressuresupport@flukecal.com with questions

'**************************************************

Function UncCalc (iT, iL, iC, iP, cTest, cConfig)

'Renaming column headers and set resolution of Meas. Unc. and GB Tol.

If cCOMPASS.CurrentTestStep = 1100 Then 'Do this at start of a new pressure/flow cycle

For i = 1 To cCOMPASS.DataCollection.Count

With cCOMPASS.DataCollection(i).DataInfoCollection("ID410008")

.Nomenclature = "Exp Unc"	'usually UserDefined8

'.Unit = "% rdg"

'.DataFormat = "0.000E-0"	'Alternate fixed way to specify resolution

'.DataFormat = "0.000"	'Alternate fixed way to specify resolution

DUTrez = cConfig.DUTPrs(1).RangeMain.ResFinal

'				cDebug.LogStatus "DUT Resolution = " & DUTrez

.DataFormat = 0.1^(DUTrez+3)	'3 additional digits of resolution for Exp Unc

'       		cDebug.LogStatus "Exp Unc Resolution = " & .DataFormat

End With

With cCOMPASS.DataCollection(i).DataInfoCollection("ID410009")

.Nomenclature = "GB Tol"	'usually UserDefined9

.DataFormat = 0.1^(DUTrez+3)	'3 additional digits of resolution for GB Tol

'       		cDebug.LogStatus "GB Tol Resolution = " & .DataFormat

End With

cCOMPASS.DataCollection(i).DataInfoCollection("ID410010").Nomenclature = "TUR"			  	'usually UserDefined10

cCOMPASS.DataCollection(i).DataInfoCollection("ID410508").Nomenclature = "Status GB"		'usually StringData8

cCOMPASS.DataCollection(i).DataInfoCollection("ID410509").Nomenclature = "Status Tol"		'usually StringData9

Next

End If

If cCompass.CurrentTestStep =  330 Then

'Using 325 is "Averaging is completed but not yet written to data file" but causes errors with multiple DUTs

'Using 330 is "Averaging is completed and logged in the data file" but doesn't update data grid

For i = 1 To cConfig.DUTPrs.Count

cDebug.LogStatus "*" 'new line between DUT's debug line in macro spy

DUTmodel = cConfig.DUTPrs(CInt(i)).RangeMain.GetParent.Model

DUTsn = cConfig.DUTPrs(CInt(i)).RangeMain.GetParent.SN

cDebug.LogStatus "DUT # " & i & ", Model " & DUTmodel & " SN " & DUTsn

dUnit = cCOMPASS.cConfig.DUTPrs(i).RangeMain.UnitFinal

'		cDebug.LogStatus "DUT Units = " & dUnit

dUnitText = cCOMPASS.cConfig.DUTPrs(i).RangeMain.UnitFinalText

'		cDebug.LogStatus "DUT Unit Text = " & dUnitText

'Get inputs for unc calc

UDut = cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).Tolerance

'		cDebug.LogStatus "DUT Tolerance = " & UDut & " " & dUnitText

DutPrs = cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).DUTPressure

'		cDebug.LogStatus "DUT Pressure = " & DutPrs & " " & dUnitText

DutRes = cConfig.DUTPrs(i).RangeMain.ResFinal

'		cDebug.LogStatus "DUT Res Rtv= " & DutRes

Res = 0.1^DutRes

cDebug.LogStatus "DUT Res = " & Res

DUTStd = cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).DUTStnDev

cDebug.LogStatus "DUT StdDev = " & DUTStd & " " & dUnitText

If DUTStd < 0 Then 'If DUT StDev <=0 set to 0

DUTStd = 0

cDebug.LogStatus "DUT StdDev (Corrected if <=0) = " & DUTStd & " " & dUnitText

End If

DUTSamples=cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).DUTCounts

cDebug.LogStatus "DUT Samples = " & DUTSamples

RefP = cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).RefPressure

RefP_Unit = cCOMPASS.cConfig.RefPrs(1).RangeMain.UnitFinal 'number of Ref Pressure unit

RefPunitTxt = cCOMPASS.cConfig.RefPrs(1).RangeMain.UnitFinalText

'		cDebug.LogStatus "Reference Pressure = " & RefP & " " & RefPunitTxt

'Below the is the normal way to get Reference Uncertainty from COMPASS Support Device setup file

URef = cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).RefTolerance

URefCP = URef 'Assign to URefCP in case we get Ref Unc by another method

'		cDebug.LogStatus "Reference Tolerance = " & UrefCP & " " & RefPunitTxt

'Uncomment the 3 following lines if you want to calculate Reference Uncertainty from the Global Macro named UncRefByRange

'	URefRange = UncRefByRange(RefP, RefPunit, RefPunitTxt)

'	URef = URefRange

'	cDebug.LogStatus "Ref Unc from Range Macro: " & URefRange & " " & RefPunitTxt

'Uncomment the 7 following lines if you want to query a 2271A, 6270A, 8270A or 8370A for Ref Pres. Uncertainty

'	cmd = "MEAS:PRES:UNC?" 'Get Reference Uncertainty from 6270A (or 2271A, 8270A or 8370A)

'	reply = cCOMPASS.cConfig.RefPrs(1).ioSendCommand(CStr(cmd),False,0)

'	cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).RefTolerance = reply

'	Uref = cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).RefTolerance

'	cDebug.LogStatus "Sent command: " & cmd

'	cDebug.LogStatus "Ref Unc from Controller: " & reply & " " & RefPunitTxt

'	cDebug.LogStatus "Ref Unc from COMPASS (not used): " & URefCP & " " & RefPunitTxt

If Uref < 0 Then 'Manual entry Ref devices report tolerance as -9999, convert to a really high number

Uref = UDut * 1000000

cDebug.LogStatus "Ref Tol (Corrected if <=0) = " & Uref & " " & RefPunitTxt

End If

cDebug.LogStatus "Reference Tolerance = " & Uref & " " & RefPunitTxt

Conf = 2 'Confidence for U1 calculation

cDebug.LogStatus "Confidence = " & Conf

k=2

cDebug.LogStatus "k = " & k

'Degrees of Freedom for S1 calculation. Comment out the lines after F=1 if you just want F=1

F = 1	'Use F=1 if number of DUT samples is 0 (zero) or less

If DUTSamples > 9 Then

F = 1.0

ElseIf DUTSamples > 7 Then

F = 1.2

ElseIf DUTSamples > 4 Then

F = 1.3

ElseIf DUTSamples > 3 Then

F = 1.7

ElseIf DUTSamples > 2 Then

F = 2.3

ElseIf DUTSamples > 0 Then

F = 7.0

End If

cDebug.LogStatus "F = " & F & ", Related to # DUT samples, lower is better"

'Calculate uncertainty

U1 = URef / Conf					'Convert Ref Unc to k=1

cDebug.LogStatus "U1 = " & U1 & " " & dUnitText & ", Uref/Confidence (Confidence usually = 2)"

S1 = (DUTStd / (DUTSamples^0.5))*F	'Calculate standard error of the mean @ k=1 and multiply by Students T factor

cDebug.LogStatus "S1 = " & S1 & " " & dUnitText & ", F * DUT Std Dev / SQRT(DUT Samples)"

S2 = (Res*0.5) / (3^0.5)			'Convert DUT resolution to k=1

cDebug.LogStatus "S2 = " & S2 & " " & dUnitText & ", Convert DUT resolution to k=2"

U2 = (((S1)^2 + (S2)^2))^0.5

cDebug.LogStatus "U2 = " & U2 & " " & dUnitText & ", RSS of S1 and S2"

Unc = (((U1)^2 + (U2)^2)^0.5)		'Combined standard uncertainty, k=1

ExpUnc = Unc * k					'Combined & expanded, k=2

cDebug.LogStatus "Unc = " & Unc & " " & dUnitText

'Test Uncertainty Ratio (TUR) with 1 digit of resolution

TUR = formatnumber(UDut/ExpUnc,1)

DUTErr = cCOMPASS.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).DUTDiff

'	cDebug.LogStatus ""	'blank line in macro spy window

'	cDebug.LogStatus "DUT Model " & DUTmodel & ", SN " & DUTsn

cDebug.LogStatus "DUT Pressure = " & DutPrs & " " & dUnitText

cDebug.LogStatus "Ref Pressure = " & RefP & " " & RefPunitTxt

cDebug.LogStatus "DUT - Ref = " & DUTErr & " " & dUnitText

'	cDebug.LogStatus "Abs(DUT-Ref) = " & Abs(DUTErr) & " " & dUnitText

cDebug.LogStatus "DUT Tolerance (no GB) = " & UDut & " " & dUnitText

cDebug.LogStatus "ExpUnc = " & ExpUnc & " " & dUnitText

'Guardband calculation and check

If (ExpUnc)^2 > (UDut)^2 Then

msgbox "Expanded Uncertainty is greater than DUT Tolerance"

GB_Tol = 0	'Set guardbanded tolerance to zero so it will fail if uncertainty is greater than DUT tolerance

Else

GB_Tol = ((UDut)^2 - (ExpUnc)^2)^0.5

cDebug.LogStatus "GB Tol = " & GB_Tol & " " & dUnitText

cDebug.LogStatus "TUR = " & TUR

End If

If GB_Tol  > Abs(DUTErr)Then

PF = "GB Pass"

Else

PF = "GB Fail"

'Uncomment the next line if you want the test to stop if the guardbanded tolerance fails

msgbox "Tolerance with Guardband failed. Press [OK], then the back button |<< if you want to repeat the last point. Press the back button until the window shows the point that you want to repeat."

End If

cDebug.LogStatus "Tolerance with guardband Pass/Fail = " & PF

If UDut > Abs(DUTErr) Then

Status = "Pass"

Else

Status = "Fail"

End If

cDebug.LogStatus "Tolerance without guardband Pass/Fail = " & Status

cCompass.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).UserDefined8 = CDbl(ExpUnc)'Expanded Uncertainty to UserDefined8

cCompass.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).UserDefined9 = CDbl(GB_Tol)'Guardbanded Tolerance to UserDefined9

cCompass.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).UserDefined10 = TUR		 'TUR to UserDefined10

cCompass.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).StringData8  = PF			'GB Pass/Fail to StringData8

cCompass.DataCollection(i).DataPointRef (CInt(iT),CInt(iL),CInt(iC),CInt(iP)).StringData9  = Status 	'Pass/Fail based on DUT-REF vs Tol to StringData9

cDebug.LogStatus "*"			'blank line in macro spy window

Next

End If

End Function

'Author: Kyle Clark

'Rev 3, 28-August-2015 added multiple DUT capability & TUR

'Rev 4, 02-Sept-2015 added units to UserDefined 1 to 3

'Rev 5, 23-Sept-2015 Changed 330 in "If cCompass.CurrentTestStep = 330" to 325 because Uncertainty

'       Parameters were not updating in data grid

'Rev 6, 07-Mar-2016 Added more comments in debug statements for more info on what S1, S2,... are.

'       Set TUR resolution to 1 digit

'Rev 7, 08-Mar-2016 Added resolution for Unc. and GB Tol. in the first few lines of code

'       Added CInt(iC) for multiple Pressure Cycle support (PresCyclIndex)

'Rev 8, 05-Apr-2016 Moved "Tolerance with Guardband failed" message box to the correct if-then-else statement

'Rev 9, 18-Jun-2018 Edited to work with multiple temperature and/or line pressure cycles. Added units to more

'       debug statements

'Rev 10, 06-May-2019 Changed to save uncertainty parameters to User8, User9, User10, StringData8 & StringData9.

'Rev 11, 29-Oct-2019 Added option to query 2271A, 6270A, 8270A or 8370A reference pressure uncertainty.

'        Added option to query Global Macro for reference pressure uncertainty.

'Rev 12, 13-Apr-2022 Changed i to 1 when getting Ref Unit and Ref Unit Text. (broke in Rev 9?)

'        The macro only supports 1 reference in a test.

'Rev 13, 20-Apr-2022 Changed from SetPrs to RefPrs

'        reply = cCOMPASS.cConfig.SetPrs(1).ioSendCommand (CStr(cmd),False,0)

'        to this in Ref Uncertainty from 6270A/2271A/8270A/8370A section of code

'        reply = cCOMPASS.cConfig.RefPrs(1).ioSendCommand (CStr(cmd),False,0)

'Rev 14, 20-May-2022 Changed directions to ensure StringData8 is in the data file and not StringData10. Added

'        CDbl assignments to some values to ensure that the values are double precision numbers and maintain

'        their values when using commas as the decimal separator.

'Rev 15, 12-Oct-2022 “325” in "If cCompass.CurrentTestStep = 325" causes the first pass of the macro to not

'work for more than one DUT. However it runs through again and works. This was changed in Rev 5 also.

'   Using 325 is "Averaging is completed but not yet written to data file" but causes errors with multiple DUTs

'   Using 330 is "Averaging is completed and logged in the data file" but doesn't update data grid