COMPASS for Flow Calibration Software

Global macro to specify Reference uncertainty by range in sccm@0C

Macro code to ignore the uncertainty values that you have entered for your references (support devices) in COMPASS for Flow software. This macro allows you to specify your reference uncertainty (in % Reading and/or % Span) according to ranges of flow. This can be used with “Measurement Uncertainty in COMPASS for Flow - Test Macro” and would be inserted after the degrees of freedom check, before the ExpUnc calculation.

Copy and Paste all of the following text into a new Global Macro.

'Alternate Reference Uncertainty Calculation/Determination - by Flow Range

'converting RefF to sccm@0C for range check

Function UncRefByRange(RefF, RefFunit, RefFunitTxt)

i=1 'use For -> Next loop for multiple DUTs

dUnit = cCOMPASS.cConfig.DUTPrs(i).RangeMain.UnitFinal 'use DUTPrs

cDebug.LogStatus "DUT Units = " & dUnit

RefF_sccm = cCOMPASS.UnitConversion(CDbl(RefF), 10, CInt(dUnit), 0) '10 is integer ID for sccm@0C

cDebug.LogStatus "RefF_sccm = " & RefF_sccm

'range specific coefs for calculation of reference uncertainty

Delta = 0	'Only use %Rdg for uncertainty unless flow is below threshold value

If RefF_sccm >= 5000 Then		'3E4-L

PctRdg = 0.2/100

ElseIf RefF_sccm >= 2000 Then	'1E4-L

PctRdg = 0.2/100

ElseIf RefF_sccm >= 200 Then	'1E3-L

PctRdg = 0.2/100

ElseIf RefF_sccm >= 20 Then		'1E2-L

PctRdg = 0.2/100

ElseIf RefF_sccm >= 2 Then 		'1E1-L %Rdg

PctRdg = 0.01/100

Else

PctRdg = 0.0			'1E1-L

Delta = 0.2			'Threshold Unc. of 0.2 sccm@0C

End If

cDebug.LogStatus "PctRdg = " & PctRdg

cDebug.LogStatus "Delta = " & Delta & " " & "sccm@0C"

'converting Delta to test unit

Delta = cCOMPASS.UnitConversion(CDbl(Delta), CInt(dUnit),1, 0)

cDebug.LogStatus "Delta = " & Delta & " " & dUnit

URef = (RefF*PctRdg) + Delta

cDebug.LogStatus "Reference Tolerance (alt) = " & Uref

End Function